| Summary of Safety and Clinical Performance | Page 1 of 42 |  |
| --- | --- | --- |
| AlboGraft™ Polyester Vascular Graft |  |  |

| 1.0 |  | Device Identification and General Information |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | i) Device trade names: AlboGraft™ Polyester Vascular Graft |  |  |  |  |  |  |  |  |
|  | ii) | Document Number: MS-0069 |  |  |  |  |  |  |  |
|  | iii) | Manufacturer’s name and address: |  |  |  |  |  |  |  |
|  |  |  | Legal manufacturer name: |  |  | LeMaitre Vascular, Inc. |  |  |  |
|  |  |  | Address: |  |  | 63 Second Avenue, Burlington, MA. 01803, USA |  |  |  |
|  | iv) SRN: US-MF-000016778 |  |  |  |  |  |  |  |  |
|  | v) | Basic UDI-DI: |  |  |  |  |  |  |  |
|  |  | • AlboGraft Woven (ATC Models): 08406631AlboGraftWV5H |  |  |  |  |  |  |  |
|  |  | • AlboGraft Knitted (AMC Models): 08406631AlboGraftKN3V |  |  |  |  |  |  |  |
|  |  | • AlboGraft Knitted Spiral Reinforced (ASC models): 08406631AlboGraftASCJL |  |  |  |  |  |  |  |
|  | vi) | Device Item Codes, Descriptions, Basic UDI |  |  |  |  |  |  |  |
|  | GTIN-14 (UDI-DI) |  |  | Item Number | Item Description |  |  |  |  |
|  | 00840663102815 |  |  | AMC1506 | Collagen-Impregnated Knitted Straight Graft 15cmx6mm |  |  |  |  |
|  |  |  |  |  | [Length x Diameter] |  |  |  |  |
|  | 00840663102822 |  |  | AMC3006 | Collagen-Impregnated Knitted Straight 30cmx6mm[LxD] |  |  |  |  |
|  | 00840663102839 |  |  | AMC4006 | Collagen-Impregnated Knitted Straight 40cmx6mm[LxD] |  |  |  |  |
|  | 00840663102846 |  |  | AMC6006 | Collagen-Impregnated Knitted Straight 60cmx6mm[LxD] |  |  |  |  |
|  | 00840663102860 |  |  | AMC4007 | Collagen-Impregnated Knitted Straight |  | 40cmx7mm [LxD] |  |  |
|  | 00840663102877 |  |  | AMC1508 | Collagen-Impregnated Knitted Straight |  | 15cmx8mm [LxD] |  |  |
|  | 00840663102884 |  |  | AMC3008 | Collagen-Impregnated Knitted Straight |  | 30cmx8mm [LxD] |  |  |
|  | 00840663102891 |  |  | AMC4008 | Collagen-Impregnated Knitted Straight |  | 40cmx8mm [LxD] |  |  |
|  | 00840663102907 |  |  | AMC6007 | Collagen-Impregnated Knitted Straight |  | 60cmx7mm [LxD] |  |  |
|  | 00840663102914 |  |  | AMC6008 | Collagen-Impregnated Knitted Straight |  | 60cmx8mm [LxD] |  |  |
|  | 00840663102921 |  |  | AMC1008 | Collagen-Impregnated Knitted Straight 100cmx8mm[LxD] |  |  |  |  |
|  | 00840663102945 |  |  | AMC3010 | Collagen-Impregnated Knitted Straight 30cmx10mm[LxD] |  |  |  |  |
|  | 00840663102952 |  |  | AMC4010 | Collagen-Impregnated Knitted Straight 40cmx10mm[LxD] |  |  |  |  |
|  | 00840663102969 |  |  | AMC6010 | Collagen-Impregnated Knitted Straight 60cmx10mm[LxD] |  |  |  |  |
|  | 00840663103010 |  |  | AMC3014 | Collagen-Impregnated Knitted Straight 30cmx14mm[LxD] |  |  |  |  |
|  | 00840663103027 |  |  | AMC1516 | Collagen-Impregnated Knitted Straight 15cmx16mm[LxD] |  |  |  |  |
|  | 00840663103034 |  |  | AMC3016 | Collagen-Impregnated Knitted Straight 30cmx16mm[LxD] |  |  |  |  |
|  | 00840663103041 |  |  | AMC1518 | Collagen-Impregnated Knitted Straight 15cmx18mm[LxD] |  |  |  |  |
|  | 00840663103058 |  |  | AMC3018 | Collagen-Impregnated Knitted Straight 30cmx18mm[LxD] |  |  |  |  |
|  | 00840663103065 |  |  | AMC1520 | Collagen-Impregnated Knitted Straight 15cmx20mm[LxD] |  |  |  |  |
|  | 00840663103072 |  |  | AMC3020 | Collagen-Impregnated Knitted Straight 30cmx20mm[LxD] |  |  |  |  |
|  | 00840663103089 |  |  | AMC1522 | Collagen-Impregnated Knitted Straight 15cmx22mm[LxD] |  |  |  |  |
|  | 00840663103096 |  |  | AMC3022 | Collagen-Impregnated Knitted Straight 30cmx22mm[LxD] |  |  |  |  |
|  | 00840663103102 |  |  | AMC1524 | Collagen-Impregnated Knitted Straight 15cmx24mm[LxD] |  |  |  |  |
|  | 00840663103119 |  |  | AMC4012 | Collagen-Impregnated Knitted Straight 40cmx12mm[LxD] |  |  |  |  |
|  | 00840663103126 |  |  | AMC3024 | Collagen-Impregnated Knitted Straight 30cmx24mm[LxD] |  |  |  |  |

SOP 35-012 rev. B

| Summary of Safety and Clinical Performance | Page 2 of 42 |  |
| --- | --- | --- |
| AlboGraft™ Polyester Vascular Graft |  |  |

| GTIN-14 (UDI-DI) | Item Number | Item Description |
| --- | --- | --- |
| 00840663103423 | AMC1207 | Collagen-impregnated Knitted Bifurcate Graft |
|  |  | 50cmx12mm,7mm [LxD1,D2] |
| 00840663107742 | AMC1206 | Collagen-impregnated Knitted Bifurcate Graft |
|  |  | 50cmx12mm,6mm [LxD1,D2] |
| 00840663103430 | AMC1407 | Collagen-impregnated Knitted Bifurcate Graft |
|  |  | 50cmx14mm,7mm [LxD1,D2] |
| 00840663103447 | AMC1408 | Collagen-impregnated Knitted Bifurcate Graft |
|  |  | 50cmx14mm,8mm [LxD1,D2] |
| 00840663103454 | AMC1608 | Collagen-impregnated Knitted Bifurcate Graft |
|  |  | 50cmx16mm,8mm [LxD1,D2] |
| 00840663103461 | AMC1609 | Collagen-impregnated Knitted Bifurcate Graft |
|  |  | 50cmx16mm,9mm [LxD1,D2] |
| 00840663103478 | AMC1809 | Collagen-impregnated Knitted Bifurcate Graft |
|  |  | 50cmx18mm,9mm [LxD1,D2] |
| 00840663103485 | AMC1810 | Collagen-impregnated Knitted Bifurcate Graft |
|  |  | 50cmx18mm,10mm [LxD1,D2] |
| 00840663103492 | AMC2010 | Collagen-impregnated Knitted Bifurcate Graft |
|  |  | 50cmx20mm,10mm [LxD1,D2] |
| 00840663103508 | AMC2011 | Collagen-impregnated Knitted Bifurcate Graft |
|  |  | 50cmx20mm,11mm [LxD1,D2] |
| 00840663103515 | AMC2211 | Collagen-impregnated Knitted Bifurcate Graft |
|  |  | 50cmx22mm,11mm [LxD1,D2] |
| 00840663103522 | AMC2412 | Collagen-impregnated Knitted Bifurcate Graft |
|  |  | 50cmx24mm,12mm [LxD1,D2] |
| 00840663104253 | AMC6012 | Collagen-Impregnated Knitted Straight 60cmx12mm[LxD] |
| 00840663104260 | AMC6014 | Collagen-Impregnated Knitted Straight 60cmx14mm[LxD] |
| 00840663104314 | AMC6024 | Collagen-Impregnated Knitted Straight 60cmx24mm[LxD] |
| 00840663107766 | AMC3007 | Collagen-Impregnated Knitted Straight 30cmx7mm[LxD] |

| GTIN-14 (UDI) | Item Number | Item Description |
| --- | --- | --- |
| 00840663103546 | ASC4006 | Collagen-impregnated Knitted Straight Graft with Removable |
|  |  | External Support 40cmx6mm[LxD] |
| 00840663103553 | ASC6006 | Collagen-impregnated Knitted Straight Graft with Removable |
|  |  | External Support 60cmx6mm[LxD] |
| 00840663103560 | ASC8006 | Collagen-impregnated Knitted Straight Graft with Removable |
|  |  | External Support 80cmx6mm[LxD] |
| 00840663103591 | ASC6007 | Collagen-impregnated Knitted Straight Graft with Removable |
|  |  | External Support 60cmx7mm[LxD] |
| 00840663103607 | ASC8007 | Collagen-impregnated Knitted Straight Graft with Removable |
|  |  | External Support 80cmx7mm[LxD] |
| 00840663103621 | ASC4008 | Collagen-impregnated Knitted Straight Graft with Removable |
|  |  | External Support 40cmx8mm[LxD] |
| 00840663103638 | ASC6008 | Collagen-impregnated Knitted Straight Graft with Removable |
|  |  | External Support 60cmx8mm[LxD] |
| 00840663103645 | ASC8008 | Collagen-impregnated Knitted Straight Graft with Removable |
|  |  | External Support 80cmx8mm[LxD] |

SOP 35-012 rev. B

| Summary of Safety and Clinical Performance | Page 3 of 42 |  |
| --- | --- | --- |
| AlboGraft™ Polyester Vascular Graft |  |  |

| GTIN-14 (UDI) | Item | Item Description |
| --- | --- | --- |
|  | Number |  |
| 00840663103669 | ATC3006 | Collagen-impregnated Woven Straight Graft 30cmx6mm[LxD] |
| 00840663103676 | ATC4006 | Collagen-impregnated Woven Straight Graft 40cmx6mm[LxD] |
| 00840663103683 | ATC6006 | Collagen-impregnated Woven Straight Graft 60cmx6mm[LxD] |
| 00840663103706 | ATC3008 | Collagen-impregnated Woven Straight Graft 30cmx8mm[LxD] |
| 00840663103713 | ATC4008 | Collagen-impregnated Woven Straight Graft 40cmx8mm[LxD] |
| 00840663103720 | ATC6008 | Collagen-impregnated Woven Straight Graft 60cmx8mm[LxD] |
| 00840663103744 | ATC3010 | Collagen-impregnated Woven Straight Graft |
|  |  | 30cmx10mm[LxD] |
| 00840663103751 | ATC4010 | Collagen-impregnated Woven Straight Graft |
|  |  | 40cmx10mm[LxD] |
| 00840663103775 | ATC1512 | Collagen-impregnated Woven Straight Graft |
|  |  | 15cmx12mm[LxD] |
| 00840663103782 | ATC3012 | Collagen-impregnated Woven Straight Graft |
|  |  | 30cmx12mm[LxD] |
| 00840663103799 | ATC4012 | Collagen-impregnated Woven Straight Graft |
|  |  | 40cmx12mm[LxD] |
| 00840663103805 | ATC1514 | Collagen-impregnated Woven Straight Graft |
|  |  | 15cmx14mm[LxD] |
| 00840663103812 | ATC3014 | Collagen-impregnated Woven Straight Graft |
|  |  | 30cmx14mm[LxD] |
| 00840663103829 | ATC1516 | Collagen-impregnated Woven Straight Graft |
|  |  | 15cmx16mm[LxD] |
| 00840663103836 | ATC3016 | Collagen-impregnated Woven Straight Graft |
|  |  | 30cmx16mm[LxD] |
| 00840663103843 | ATC1518 | Collagen-impregnated Woven Straight Graft |
|  |  | 15cmx18mm[LxD] |
| 00840663103850 | ATC3018 | Collagen-impregnated Woven Straight Graft |
|  |  | 30cmx18mm[LxD] |
| 00840663103867 | ATC1520 | Collagen-impregnated Woven Straight Graft |
|  |  | 15cmx20mm[LxD] |
| 00840663103874 | ATC3020 | Collagen-impregnated Woven Straight Graft |
|  |  | 30cmx20mm[LxD] |
| 00840663103881 | ATC1522 | Collagen-impregnated Woven Straight Graft |
|  |  | 15cmx22mm[LxD] |
| 00840663103898 | ATC3022 | Collagen-impregnated Woven Straight Graft |
|  |  | 30cmx22mm[LxD] |
| 00840663103911 | ATC1524 | Collagen-impregnated Woven Straight Graft |
|  |  | 15cmx24mm[LxD] |
| 00840663103928 | ATC3024 | Collagen-impregnated Woven Straight Graft |
|  |  | 30cmx24mm[LxD] |
| 00840663103942 | ATC1526 | Collagen-impregnated Woven Straight Graft |
|  |  | 15cmx26mm[LxD] |
| 00840663103959 | ATC3026 | Collagen-impregnated Woven Straight Graft |
|  |  | 30cmx26mm[LxD] |
| 00840663103973 | ATC1528 | Collagen-impregnated Woven Straight Graft |
|  |  | 15cmx28mm[LxD] |
| 00840663103980 | ATC3028 | Collagen-impregnated Woven Straight Graft |
|  |  | 30cmx28mm[LxD] |
| 00840663104000 | ATC1530 | Collagen-impregnated Woven Straight Graft |
|  |  | 15cmx30mm[LxD] |
| 00840663104017 | ATC3030 | Collagen-impregnated Woven Straight Graft |
|  |  | 30cmx30mm[LxD] |
| 00840663104031 | ATC1532 | Collagen-impregnated Woven Straight Graft |
|  |  | 15cmx32mm[LxD] |
| 00840663104048 | ATC3032 | Collagen-impregnated Woven Straight Graft |
|  |  | 30cmx32mm[LxD] |
| 00840663104062 | ATC1534 | Collagen-impregnated Woven Straight Graft |
|  |  | 15cmx34mm[LxD] |

SOP 35-012 rev. B

| Summary of Safety and Clinical Performance | Page 4 of 42 |  |
| --- | --- | --- |
| AlboGraft™ Polyester Vascular Graft |  |  |

| GTIN-14 (UDI) | Item | Item Description |
| --- | --- | --- |
|  | Number |  |
| 00840663104079 | ATC3034 | Collagen-impregnated Woven Straight Graft |
|  |  | 30cmx34mm[LxD] |
| 00840663104093 | ATC1538 | Collagen-impregnated Woven Straight Graft |
|  |  | 15cmx38mm[LxD] |
| 00840663104109 | ATC3038 | Collagen-impregnated Woven Straight Graft |
|  |  | 30cmx38mm[LxD] |
| 00840663104147 | ATC1207 | Collagen-impregnated Woven Bifurcated Graft |
|  |  | 50cmx12mm[LxD1,D2] |
| 00840663110384 | ATC1206 | Collagen-impregnated Woven Bifurcated Graft 50mm x |
|  |  | 12mm,6mm [LxD1,D2] |
| 00840663104154 | ATC1407 | Collagen-impregnated Woven Bifurcated Graft |
|  |  | 50cmx14mm,7mm [LxD1,D2] |
| 00840663104161 | ATC1408 | Collagen-impregnated Woven Bifurcated Graft |
|  |  | 50cmx14mm,8mm [LxD1,D2] |
| 00840663104178 | ATC1608 | Collagen-impregnated Woven Bifurcated Graft |
|  |  | 50cmx16mm,8mm [LxD1,D2] |
| 00840663104185 | ATC1609 | Collagen-impregnated Woven Bifurcated Graft |
|  |  | 50cmx16mm,9mm [LxD1,D2] |
| 00840663104192 | ATC1809 | Collagen-impregnated Woven Bifurcated Graft |
|  |  | 50cmx18mm,9mm [LxD1,D2] |
| 00840663104208 | ATC1810 | Collagen-impregnated Woven Bifurcated Graft |
|  |  | 50cmx18mm,10mm [LxD1,D2] |
| 00840663104215 | ATC2010 | Collagen-impregnated Woven Bifurcated Graft |
|  |  | 50cmx20mm,10mm [LxD1,D2] |
| 00840663104239 | ATC2211 | Collagen-impregnated Woven Bifurcated Graft |
|  |  | 50cmx22mm,11mm [LxD1,D2] |
| 00840663104246 | ATC2412 | Collagen-impregnated Woven Bifurcated Graft |
|  |  | 50cmx24mm,12mm [LxD1,D2] |
| 00840663104437 | ATC4007 | Collagen-impregnated Woven Straight Graft 40cmx7mm[LxD] |
| 00840663104444 | ATC6012 | Collagen-impregnated Woven Straight Graft |
|  |  | 60cmx12mm[LxD] |
| 00840663104451 | ATC6014 | Collagen-impregnated Woven Straight Graft |
|  |  | 60cmx14mm[LxD] |
| 00840663104468 | ATC6016 | Collagen-impregnated Woven Straight Graft |
|  |  | 60cmx16mm[LxD] |
| 00840663104475 | ATC6018 | Collagen-impregnated Woven Straight Graft |
|  |  | 60cmx18mm[LxD] |
| 00840663104482 | ATC6020 | Collagen-impregnated Woven Straight Graft |
|  |  | 60cmx20mm[LxD] |
| 00840663104499 | ATC6022 | Collagen-impregnated Woven Straight Graft |
|  |  | 60cmx22mm[LxD] |
| 00840663104505 | ATC6024 | Collagen-impregnated Woven Straight Graft |
|  |  | 60cmx24mm[LxD] |
| 00840663104512 | ATC6026 | Collagen-impregnated Woven Straight Graft |
|  |  | 60cmx26mm[LxD] |
| 00840663104529 | ATC6028 | Collagen-impregnated Woven Straight Graft |
|  |  | 60cmx28mm[LxD] |
| 00840663104536 | ATC6030 | Collagen-impregnated Woven Straight Graft |
|  |  | 60cmx30mm[LxD] |
| 00840663106677 | ATC3036 | Collagen-impregnated Woven Straight Graft 30cmx36mm |
|  |  | [LxD] |
| 00840663107407 | ATC6007 | Collagen-impregnated Woven Straight Graft 60cmx7mm[LxD] |
| 00840663110353 | ATC6032 | Collagen-impregnated Woven Straight Graft 60cm x 32mm |
|  |  | [LxD] |
| 00840663110377 | ATC6038 | Collagen-impregnated Woven Straight Graft 60cm x 38mm |
|  |  | [LxD] |

SOP 35-012 rev. B

| Summary of Safety and Clinical Performance | Page 5 of 42 |  |
| --- | --- | --- |
| AlboGraft™ Polyester Vascular Graft |  |  |

vii) Medical device nomenclature description

GMDN Code / Description: 35281 / Synthetic Vascular Graft

UMDNS Code / Description: 13-177 / Prostheses, Blood Vessel, Artificial

EMDN Code / Description: P07010201 / Vascular prosthesis, dacron

| viii) |  | Class of device |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Manufacture Name |  |  |  | MDR Classification |  |  |  | Rule |  |  |
|  |  | AlboGraft Polyester Vascular Graft |  |  |  | III Implantable |  |  |  | 18 |  |  |
| ix) Year when the first certificate (CE) was issued covering the device |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  | Device Name |  | Date of Initial |  |  | Date of 510(k) |  |  |  |  |  |
|  |  |  |  | CE Mark |  |  |  |  |  |  |  |  |
|  |  | AlboGraft™ Polyester |  | 15 April 2011 |  |  | 14 January 2010 (K093231) |  |  |  |  |  |
|  |  | Vascular Graft |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  | 19 January 2011 (K103080) |  |  |  |  |  |
| x) Authorised representative if applicable; name and the SRN |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  |  | EU Authorized |  | LeMaitre Vascular GmbH |  |  |  |  |  |  |  |
|  |  |  | Representative: |  | Otto-Volger-Str. 5 a/b |  |  |  |  |  |  |  |
|  |  |  |  |  | 65843, Sulzbach/Ts |  |  |  |  |  |  |  |
|  |  |  |  |  | Germany |  |  |  |  |  |  |  |
|  |  |  | SRN: |  | DE-AR-000013539 |  |  |  |  |  |  |  |

xi)	NB’s name (the NB that will validate the SSCP) and the NB’s single identification number

BSI Group The Netherlands B.V. Identification Number: 2797

Say Building, John M. Keynesplein 9, 1066 EP Amsterdam, Netherlands

- Intended use of the device

  - Intended purpose: The AlboGraft Vascular Grafts are designed for the purpose of systemic vascular repair including replacement or bypass procedures in aneurysmal and occlusive disease of the arteries including the thoracic aorta and for femoral popliteal reconstruction in the treatment of abdominal aortic aneurysm, thoracic aortic aneurysm, and peripheral artery disease.

  - Indication(s) and target population(s)

    - Indication:

      - The AlboGraft Knitted and Woven Vascular Grafts are indicated for use in the replacement or repair of arteries affected with aneurismal or occlusive disease, such as abdominal aortic aneurysm, thoracic aortic aneurysm, and peripheral artery

SOP 35-012 rev. B

| Summary of Safety and Clinical Performance | Page 6 of 42 |  |
| --- | --- | --- |
| AlboGraft™ Polyester Vascular Graft |  |  |

disease (e.g., TASC C or D type lesions) involving the iliac, femoral, and popliteal arteries.

      - The AlboGraft Vascular Graft (ASC models only) are indicated in extra-anatomic reconstructions and reconstructions requiring enhanced resistance to kinking and compression, such as femoro-popliteal bypass.

      - Target Population: Adults of any gender or ethnicity in need of systemic vascular repair including replacement or bypass procedures in aneurysmal and occlusive disease of the arteries including the thoracic aorta and for femoral popliteal reconstruction.

  - Contraindications and/or limitations

    - AlboGraft Vascular Grafts are contraindicated for use in coronary arteries.

    - AlboGraft Vascular Grafts are contraindicated in patients with known or suspected hypersensitivity to bovine collagen.

- Device Description

  - Description of the device

AlboGraft® Polyester Vascular Grafts are made of synthetic material and designed to replace sections of damaged or malfunctioning arteries. They are made of polyester (polyethylene terephthalate, PET) thread woven or knitted into a continuous tube. In response to a range of surgical indications, AlboGraft Vascular Grafts are oﬀered in two designs: double velour knitted fabric and double velour woven fabric. The knitted grafts are designed with a run-proof structure to reduce the risk of fraying or wearing down at their ends. The velour grafts have low-profile loops on their endoluminal surface to avoid any lumen reduction, and high-profile loops on their outer surface to promote graft anchoring into the surrounding tissues. All AMC and ATC AlboGraft grafts are crimped in parallel rings so that their tubular shape is maintained without kinking.

AlboGraft Vascular Grafts are available with removable external spiral reinforcement (ASC models) made of a radiopaque, biocompatible thread with a polypropylene core and polyethylene coating, allowing for easy identification of the prosthesis with x-ray. The external spiral reinforcement is removable, facilitating the creation of anastomoses to the vessel.

Note: AlboGraft Vascular Grafts with Removable External Spiral (ASC Models) are not available for sale in Canada.

Guidelines on the outer surface of the graft facilitate orientation during implantation.

The AlboGraft Vascular Grafts impregnated with collagen reduce the permeability to mitigate the problem of hemorrhaging during implantation so that no preclotting is necessary. The process of impregnation using bovine collagen maintains both the original structure of the material, and the structural characteristics of the prosthesis, i.e. flexibility and softness. The

SOP 35-012 rev. B

| Summary of Safety and Clinical Performance | Page 7 of 42 |  |
| --- | --- | --- |
| AlboGraft™ Polyester Vascular Graft |  |  |

| collagen is cross-linked by formaldehyde steaming to ensure a balanced reabsorption rate by |  |
| --- | --- |
| the host tissues. |  |
| Image | Device name |
|  | LeMaitre AlboGraft™ Polyester Vascular Graft |
|  | Configurations (straight, bifurcated, external |
|  | supported) |

- Reference to previous generations: The product is a mature product currently on the market for a well-established intended use. It has been developed by incremental changes and is based on the Hemashield Microvel Double Velour Knitted and Woven Vascular Graft. The AlboGraft was previously manufactured by Biomateriali S. r. l., a subsidiary of LeMaitre Vascular, Inc. in Brindisi, Italy. LeMaitre Vascular has transferred the manufacturing from Italy to Burlington MA.

- There are no novel design features, indications, claims, or target populations for the subject device compared to the competitor devices that impact safety and performance, although minor changes have been made to the device to provide incremental benefits to the user/patients. A primary difference is the option of grafts with and without collagen impregnation for the AlboGraft Polyester Vascular Grafts.

- Description of any accessories which are intended to be used in combination with the device: No accessories are supplied with this device.

- Description of any other devices and products which are intended to be used in combination with the device: No other devices or products are intended to be used in combination with this device.

SOP 35-012 rev. B

| Summary of Safety and Clinical Performance | Page 8 of 42 |  |
| --- | --- | --- |
| AlboGraft™ Polyester Vascular Graft |  |  |

- Risks and Warnings

  - Residual risks and undesirable effects

    - Residual risk evaluation is conducted as part of our FMEAs and risk management procedure. We have concluded that the benefits outweigh any residual risks and that the risk has been reduced as far as possible.

-	Potential Complications:

| Adverse Event | Rate | Source from CER |
| --- | --- | --- |
| Stroke | 6% | Hsu, #37 |
| Paraplegia | 0% | Post market study 2009 |
| Paraparesis | 1% | Biomateriali, 2008 |
| Myocardial Infarction | 1% | Biomateriali, 2008 |
| Renal Dysfunction | 2.8% | Lamelas, #38 |
| Ischemia | 5% | Almasri, #1 |
| Embolism | - | No reported occurrence |
| Thrombosis | 1.3% | Kim, #35 |
| Bleeding | 2% | Biomateriali 2009 |
| Graft Infection | - | No reported occurrence |
| Wound Infection | 1% | Biomaterilai, 2008 |
| Aneurysm | - | No reported occurrence |
| Pneumonia | - | No reported occurrence |
| Amputation | 1.6% | Biomateriali, 2010 |
| Death | <6.6% | Tamura, #40 |
| Graft Dilation | - | No reported occurrence |
| Graft disintegration | - | No reported occurrence |
| Graft stenosis | - | No reported occurrence |
| Pseudoaneurysm | - | No reported occurrence |
| Intraluminal graft thrombus | - | No reported occurrence |
| Perigraft air | - | No reported occurrence |
| Respiratory failure | 5% | Hsu, #37 |
| Arterial fibrillation | 27.7% | Lamelas, #38 |
| Chylothorax | 4.7% | Rajbanshi, #39 |
| Temporary psychotic | 8% | Hsu, #37 |
| syndrome |  |  |
| Hemiparesis | 3.9% | Biomateriali, 2009 |
| Visceral ischemia | 3.9% | Biomateriali, 2009 |
| Gluteal necrosis | 1% | Biomaterilai, 2008 |
| Ileus | 2.9% | Biomaterilai, 2008 |
| Acute kidney failure | 1.3% | Biomaterilai, 2009 |
| Compartment syndrome | 1.3% | Biomaterilai, 2009 |
| Paravalvular leak | 1.3% | Biomaterilai, 2008 |
| Descending aorta dissection | 1.4% | Biomaterilai, 2009 |
| Deep vein thrombosis | 1.4% | Biomaterilai, 2009 |
| Sternum instability | 1.4 | Biomaterilai, 2009 |
| Bypass occlusions | 1.6% | Biomaterilai, 2010 |

SOP 35-012 rev. B

| Summary of Safety and Clinical Performance | Page 9 of 42 |  |
| --- | --- | --- |
| AlboGraft™ Polyester Vascular Graft |  |  |

| Wound hematoma | 1.5% | Biomaterilai, 2010 |
| --- | --- | --- |
| Groin Seroma | 1.5% | Biomaterilai, 2010 |
| Urinary tract infection | 1.5% | Biomaterilai, 2010 |
| Angina abdominalis | 1.6% | Biomaterilai, 2010 |

- Warnings and precautions

  - Do not use a prosthesis if the container and/or seal has been opened or damaged, or if the period of sterility has expired.

  - The collagen-impregnated graft must never be re-sterilized.

  - Grafts contaminated with blood during the preceding procedures must not be re-used or re-sterilized.

  - The vascular grafts must be handled so as to avoid contact with extraneous particles which, if they adhere to the graft wall, may generate emboli or undesirable interactions with the blood.

  - Furthermore, surgical gloves used to handle grafts should not contain powders, preservatives or lubricants.

  - Avoid overstretching the graft; gently expand the graft to smooth the folds.

  - Avoid damaging the graft when handling, use atraumatic clamps and appropriate instruments (e.g. vascular clamps). Avoid using these instruments with undue force, otherwise the collagen coating or fabric will be damaged.

  - Atraumatic needles are recommended.

  - Low temperature ophthalmic cautery (≤ 704° C/1300° F) is recommended for cutting woven grafts to avoid fraying.

  - AlboGraft Removable Spiral Reinforcement Prosthesis (ASC Models) *: Avoid clamping the graft on its reinforced area.

  - AlboGraft Removable Spiral Reinforcement Prosthesis (ASC Models) *: Gently remove the support spiral, otherwise the collagen film will be damaged.

  - Care should be taken to ligate and/or cauterize lymphatics in the groin to minimize the occurrence of seroma formation and lymphatic collection after aorto-femoral or femoropopliteal reconstruction.

  - These prostheses should not be implanted in patients who exhibit sensitivity to polyester or materials of bovine origin.

- Other relevant aspects of safety, including a summary of any field safety corrective action (FSCA including FSN) if applicable.

    - During 01 January 2019 to 31 May 2024, there were a total of 184 complaints associated with the subject device and a total of 47,011 devices sold, resulting in an overall cumulative complaint rate of 0.391%. The table below provides the complaint rate for each year. The rates of each complaint are in alignment with the frequency in the risk management documentation.

SOP 35-012 rev. B

| Summary of Safety and Clinical Performance | Page 10 of 42 |  |
| --- | --- | --- |
| AlboGraft™ Polyester Vascular Graft |  |  |

Complaints rates per year

| Complaints by Region / Year | 2019 | 2020 | 2021 | 2022 | 2023 | 2024* | Total |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total Complaints | 9 | 8 | 2 | 24 | 57 | 84 | 184 |  |
| Total Sales | 9,252 | 7,459 | 7,058 | 8,761 | 10,372 | 4,109 | 47,011 |  |
| Total Complaint Rate | 0.097% | 0.107% | 0.028% | 0.274% | 0.550% | 2.044% | 0.391% |  |
| Europe | 2019 | 2020 | 2021 | 2022 | 2023 | 2024* | Total |  |
| Complaints | 7 | 8 | 2 | 15 | 42 | 83 | 157 |  |
| Sales | 7,837 | 5,748 | 5,343 | 6,258 | 7,405 | 3,008 | 35,599 |  |
| Rate (complaints/sales) | 0.089% | 0.139% | 0.037% | 0.240% | 0.567% | 2.759% | 0.441% |  |
| Americas | 2019 | 2020 | 2021 | 2022 | 2023 | 2024* | Total |  |
| Complaints | 2 | 0 | 0 | 9 | 1 | 0 | 12 |  |
| Sales | 1,414 | 1,707 | 1,707 | 2,474 | 2,837 | 918 | 11,057 |  |
| Rate (complaints/sales) | 0.141% | 0.000% | 0.000% | 0.364% | 0.035% | 0.000% | 0.109% |  |
| Asia | 2019 | 2020 | 2021 | 2022 | 2023 | 2024* | Total |  |
| Complaints | 0 | 0 | 0 | 0 | 14 | 1 | 15 |  |
| Sales | 1 | 4 | 8 | 29 | 130 | 183 | 355 |  |
| Rate (complaints/sales) | - | - | 0.000% | 0.000% | 10.769% | 0.546% | 4.225% |  |

*Through May

Complaint rates by model type per year

| Model ID | Year | # | # Devices | Complaint | # Adverse | Adverse |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Complaints | sold | rate | events | event rate |  |
| AMC: Knitted | 2019 | 5 | 2848 | 0.176% | 0 | 0.000% |  |
| Straight |  |  |  |  |  |  |  |
|  | 2020 | 3 | 2313 | 0.130% | 3 | 0.130% |  |
|  | 2021 | 0 | 2251 | 0.000% | 0 | 0.000% |  |
|  | 2022 | 9 | 2531 | 0.356% | 1 | 0.040% |  |
|  | 2023 | 21 | 2782 | 0.755% | 0 | 0.000% |  |
|  | 2024 | 29 | 1156 | 2.509% | 0 | 0.000% |  |
|  | Total | 67 | 13881 | 0.483% | 4 | 0.029% |  |
| AMC: Knitted | 2019 | 3 | 2224 | 0.135% | 1 | 0.045% |  |
| Bifurcated |  |  |  |  |  |  |  |
|  | 2020 | 1 | 1796 | 0.056% | 1 | 0.056% |  |
|  | 2021 | 2 | 1721 | 0.116% | 2 | 0.116% |  |
|  | 2022 | 12 | 2314 | 0.519% | 1 | 0.043% |  |
|  | 2023 | 2 | 2636 | 0.076% | 0 | 0.000% |  |
|  | 2024 | 17 | 915 | 1.858% | 2 | 0.219% |  |
|  | Total | 37 | 11606 | 0.319% | 7 | 0.060% |  |

SOP 35-012 rev. B

| Summary of Safety and Clinical Performance | Page 11 of 42 |  |
| --- | --- | --- |
| AlboGraft™ Polyester Vascular Graft |  |  |

| Model ID | Year | # | # Devices | Complaint | # Adverse | Adverse |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Complaints | sold | rate | events | event rate |  |
| ASC: Knitted | 2019 | 0 | 1040 | 0.000% | 0 | 0.000% |  |
| Straight with |  |  |  |  |  |  |  |
|  | 2020 | 2 | 858 | 0.233% | 0 | 0.000% |  |
| Removable |  |  |  |  |  |  |  |
|  | 2021 | 0 | 733 | 0.000% | 0 | 0.000% |  |
| External Spiral |  |  |  |  |  |  |  |
| Support | 2022 | 0 | 813 | 0.000% | 0 | 0.000% |  |
|  | 2023 | 18 | 814 | 2.211% | 0 | 0.000% |  |
|  | 2024 | 8 | 426 | 1.878% | 0 | 0.000% |  |
|  | Total | 28 | 4684 | 0.598% | 0 | 0.000% |  |
| ATC: Woven | 2019 | 1 | 2788 | 0.036% | 0 | 0.000% |  |
| Straight |  |  |  |  |  |  |  |
|  | 2020 | 2 | 2232 | 0.090% | 0 | 0.000% |  |
|  | 2021 | 0 | 2027 | 0.000% | 0 | 0.000% |  |
|  | 2022 | 3 | 2709 | 0.111% | 0 | 0.000% |  |
|  | 2023 | 8 | 3617 | 0.221% | 0 | 0.000% |  |
|  | 2024 | 28 | 1397 | 2.004% | 0 | 0.000% |  |
|  | Total | 42 | 14770 | 0.284% | 0 | 0.000% |  |
| ATC: Woven | 2019 | 0 | 352 | 0.000% | 0 | 0.000% |  |
| Bifurcated |  |  |  |  |  |  |  |
|  | 2020 | 0 | 260 | 0.000% | 0 | 0.000% |  |
|  | 2021 | 0 | 326 | 0.000% | 0 | 0.000% |  |
|  | 2022 | 0 | 394 | 0.000% | 0 | 0.000% |  |
|  | 2023 | 6 | 523 | 1.147% | 0 | 0.000% |  |
|  | 2024 | 2 | 215 | 0.930% | 0 | 0.000% |  |
|  | Total | 8 | 2070 | 0.386% | 0 | 0.000% |  |
| Unknown | 2019 | 0 | 352 | 0.000% | 0 | 0.000% |  |
|  | 2020 | 0 | 260 | 0.000% | 0 | 0.000% |  |
|  | 2021 | 0 | 326 | 0.000% | 0 | 0.000% |  |
|  | 2022 | 0 | 394 | 0.000% | 0 | 0.000% |  |
|  | 2023 | 2 | 523 | 0.382% | 2 | 0.382% |  |
|  | 2024 | 0 | 215 | 0.000% | 0 | 0.000% |  |
|  | Total | 2 | 2070 | 0.097% | 2 | 0.097% |  |

Complaints by type per year

| Complaint Category | 2019 | 2020 | 2021 | 2022 | 2023 | 2024* | Total | Rate |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Red temperature indicator** | 0 | 0 | 0 | 0 | 42 | 29 | 71 | 0.151 |
| Defective temperature indicator** | 0 | 0 | 0 | 1 | 7 | 46 | 54 | 0.115 |
| Packaging issue | 8 | 2 | 0 | 10 | 4 | 0 | 24 | 0.051 |
| Shipping damage | 0 | 0 | 0 | 9 | 0 | 1 | 10 | 0.021 |
| Difficult to remove the monofilament | 0 | 0 | 0 | 0 | 1 | 4 | 5 | 0.011 |
| Hole in the graft | 0 | 1 | 2 | 1 | 0 | 0 | 4 | 0.009 |
| Contamination by the user | 0 | 0 | 0 | 1 | 2 | 0 | 3 | 0.006 |

SOP 35-012 rev. B

| Summary of Safety and Clinical Performance | Page 12 of 42 |  |
| --- | --- | --- |
| AlboGraft™ Polyester Vascular Graft |  |  |

| Complaint Category | 2019 | 2020 | 2021 | 2022 | 2023 | 2024* | Total | Rate |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Barcode error | 0 | 0 | 0 | 2 | 0 | 0 | 2 | 0.004 |
| Graft thrombosis/occlusion | 0 | 2 | 0 | 0 | 0 | 0 | 2 | 0.004 |
| Leaking | 1 | 0 | 0 | 0 | 0 | 1 | 2 | 0.004 |
| Service issue - incorrect product delivered | 0 | 0 | 0 | 0 | 0 | 2 | 2 | 0.004 |

| Collagen fell off | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0.002 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Difficult to remove tyvek lid | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0.002 |
| Graft rupture | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0.002 |
| Monofilament brittle | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0.002 |
| User error | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0.002 |

*Through May ** CAPA 2024-010 opened to address higher rates

- From 01 January 2019 to 31 May 2024 there were 4 CAPAs opened, more details in the tables below.

CAPA summary

| CAPA # | Description | Date | Date | Status |  |
| --- | --- | --- | --- | --- | --- |
|  |  | Initiated | Closed |  |  |
| CAPA | Reason CAPA initiated: Product complaints have been | 14-Apr-23 | N/A | In process |  |
| 2023-016 | received indicating that the handling instructions for the |  |  |  |  |
| Memo | AlboGraft product is not clear enough and it can lead to |  |  |  |  |
| 2023- | users accidentally introducing the non-sterile outer |  |  |  |  |
| 0032 | package into the sterile field. |  |  |  |  |
|  | Corrective action taken: |  |  |  |  |
|  | Preventive Action-Phase 1: adhere a new sticker to the |  |  |  |  |
|  | AlboGraft outer Tyvek lid stating “Exterior of this |  |  |  |  |
|  | Package is Not Sterile” along with the Sterile EO and |  |  |  |  |
|  | Double Sterile Barrier symbols per ISO15223-1:2021 |  |  |  |  |
|  | Preventive Action-Phase 2: The artwork on the |  |  |  |  |
|  | AlboGraft outer Tyvek lid will be updated to integrate |  |  |  |  |
|  | the sticker artwork from Phase 1, with Sterile EO and |  |  |  |  |
|  | Double Sterile Barrier symbols reallocated to the |  |  |  |  |
|  | corresponding symbols area of the label, so that the |  |  |  |  |
|  | application of the sticker can be eliminated; and, the |  |  |  |  |
|  | outer foil pouch artwork layout will be reorganized / |  |  |  |  |
|  | streamlined. |  |  |  |  |
| CAPA | Reason CAPA initiated: AlboGraft labels not matching | 19-Aug-19 | 23-Mar-21 | Closed |  |
| 2019-055 | on all of the packaging. |  |  |  |  |
|  | Corrective action taken: Updated AlboGraft Mis to |  |  |  |  |
|  | include procedure for label verification between |  |  |  |  |
|  | packaging levels. |  |  |  |  |
| CAPA | Reason CAPA initiated: | 28-Mar-24 | N/A | In process |  |
| 2024-006 | 1. Poor quality of the temp. indicators from the supplier. |  |  |  |  |

SOP 35-012 rev. B

| Summary of Safety and Clinical Performance | Page 13 of 42 |  |
| --- | --- | --- |
| AlboGraft™ Polyester Vascular Graft |  |  |

| CAPA # | Description |  | Date | Date | Status |
| --- | --- | --- | --- | --- | --- |
|  |  |  | Initiated | Closed |  |
|  | 2. | Storing the indicators in a freezer, not a fridge |  |  |  |
|  | 3. | Tightly storing AlboGraft boxes in white cardboard |  |  |  |
|  | boxes in Finished Goods inventory. Removing each box |  |  |  |  |
|  | from this tightly packed box, initiates friction and |  |  |  |  |
|  | damages the temp. indicator. |  |  |  |  |
|  | 4. | Inadequate packaging instructions in SOP48-001 |  |  |  |
|  | (Shipping Temperature Sensitive Product Procedure) |  |  |  |  |
|  | and SOPlS-004 (Shipping Product Procedure) - missing |  |  |  |  |
|  | AlboGraft bulk shipping process |  |  |  |  |
|  | Corrective action taken: On-going |  |  |  |  |
| CAPA | Reason CAPA initiated: Investigation on-going |  | 17-May-24 | N/A | In process |
| 2024-010 |  |  |  |  |  |

- FSCA: There were 3 recall notifications sent out by LeMaitre for the AlboGraft product family during the reporting period of 01 January 2019 to 31 May 2024.

| Date Initiated | Recall # | Region | Products | Description |  |
| --- | --- | --- | --- | --- | --- |
|  | /FSCA |  | Involved |  |  |
|  | number |  |  |  |  |
| 15 February | Urgent field | Spain market | Model | LeMaitre is withdrawing and |  |
|  | safety notice |  |  |  |  |
| 2022 |  |  | Numbers: | exchanging AlboGraft devices |  |
|  | BfArM |  |  |  |  |
|  |  |  | AMC1407 | without a CE mark due to a change |  |
|  | Reference: |  |  |  |  |
|  |  |  | AMC1408 | in regulatory status. |  |
|  | 04068/22 |  |  |  |  |
|  |  |  | AMC1516 | BACKGROUND: In 2020, AEMPS |  |
|  |  |  | AMC1518 | granted LeMaitre permission |  |
|  |  |  | AMC1520 | (derogation) to supply AlboGraft |  |
|  |  |  | AMC1522 | devices on the Spain market, |  |
|  |  |  | AMC1608 | without a CE mark. This derogation |  |
|  |  |  | AMC1809 | was requested because our previous |  |
|  |  |  | AMC2010 | notified body had stopped providing |  |
|  |  |  | AMC2211 | CE marking services and our new |  |
|  |  |  | AMC3006 | notified body had not completed |  |
|  |  |  | AMC3008 | their on-boarding review of the |  |
|  |  |  | AMC3010 | technical documentation. We are |  |
|  |  |  | AMC3014 | grateful to AEMPS for allowing us |  |
|  |  |  | AMC3016 | to serve our customers while we |  |
|  |  |  | AMC3018 | worked to gain approval from our |  |
|  |  |  | AMC3020 | new notified body. AEMPS has |  |
|  |  |  | AMC3022 | received a copy of this letter. |  |
|  |  |  | AMC4006 | LeMaitre has now received the CE |  |
|  |  |  | AMC4007 | mark approval from our new |  |
|  |  |  | AMC4008 | notified body for AlboGraft and we |  |

SOP 35-012 rev. B

| Summary of Safety and Clinical Performance | Page 14 of 42 |  |
| --- | --- | --- |
| AlboGraft™ Polyester Vascular Graft |  |  |

| Date Initiated | Recall # | Region | Products | Description |  |
| --- | --- | --- | --- | --- | --- |
|  | /FSCA |  | Involved |  |  |
|  | number |  |  |  |  |
|  |  |  | AMC6006 | now have adequate capacity to |  |
|  |  |  | AMC6007 | serve all of our European customers |  |
|  |  |  | AMC6008 | with CE marked product. Although |  |
|  |  |  | AMC6010 | there is no safety risk with any of |  |
|  |  |  | ASC4006 | the derogated devices, we are |  |
|  |  |  | ASC4008 | required to withdraw any remaining |  |
|  |  |  | ASC6006 | devices as a condition of the |  |
|  |  |  | ASC6008 | derogation |  |
|  |  |  | ASC8006 |  |  |
|  |  |  | ASC8008 |  |  |
|  |  |  | ATC1207 |  |  |
| 18 February | Urgent field | Dutch market | AMC1407 | LeMaitre is withdrawing and |  |
|  | safety notice |  |  |  |  |
| 2022 |  |  | AMC1608 | exchanging certain AlboGraft |  |
|  | BfArM |  |  |  |  |
|  |  |  | AMC1809 | devices due to a change in |  |
|  | Reference: |  |  |  |  |
|  |  |  | AMC2211 | regulatory status. BACKGROUND: |  |
|  | 03253A/22 |  |  |  |  |
|  |  |  | AMC3016 | In 2020, Ministerie van |  |
|  |  |  | AMC3018 | Volksgezondheid, Welzijn en Sport |  |
|  |  |  | AMC3020 | granted LeMaitre permission |  |
|  |  |  | AMC6006 | (derogation) to supply AlboGraft |  |
|  |  |  | ASC4006 | devices on the Dutch market, |  |
|  |  |  | ASC4008 | without a CE mark. This derogation |  |
|  |  |  | ASC8006 | was requested because our previous |  |
|  |  |  | ASC8008 | notified body had stopped providing |  |
|  |  |  | ATC1526 | CE marking services and our new |  |
|  |  |  | ATC1528 | notified body had not completed |  |
|  |  |  | ATC1530 | their on-boarding review of the |  |
|  |  |  | ATC3006 | technical documentation. We are |  |
|  |  |  | ATC3008 | grateful to Ministerie van |  |
|  |  |  | ATC3028 | Volksgezondheid, Welzijn en Sport |  |
|  |  |  | ATC3030 | for allowing us to serve our |  |
|  |  |  | ATC3032 | customers while we worked to gain |  |
|  |  |  | ATC4006 | approval from our new notified |  |
|  |  |  |  | body. They have received a copy of |  |
|  |  |  |  | this letter. LeMaitre has now |  |
|  |  |  |  | received the CE mark approval |  |
|  |  |  |  | from our new notified body for |  |
|  |  |  |  | AlboGraft and we now have |  |
|  |  |  |  | adequate capacity to serve all of our |  |
|  |  |  |  | European customers with CE |  |
|  |  |  |  | marked product. Although there is |  |
|  |  |  |  | no safety risk with any of the |  |
|  |  |  |  | derogated devices, we are required |  |

SOP 35-012 rev. B

|  |  |  | Summary of Safety and Clinical Performance |  |  | Page 15 of 42 |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | AlboGraft™ Polyester Vascular Graft |  |  |  |  |  |
|  | Date Initiated | Recall # | Region | Products | Description |  |  |  |
|  |  | /FSCA |  | Involved |  |  |  |  |
|  |  | number |  |  |  |  |  |  |
|  |  |  |  |  | to withdraw any remaining devices |  |  |  |
|  |  |  |  |  | as a condition of the derogation. |  |  |  |
|  | 01 March | MHRA | UK | AMC1008 | LeMaitre is withdrawing and |  |  |  |
|  |  | reference |  |  |  |  |  |  |
|  | 2022 |  |  | AMC1206 | exchanging AlboGraft devices |  |  |  |
|  |  | number: |  |  |  |  |  |  |
|  |  |  |  | AMC1207 | without a CE mark due to a change |  |  |  |
|  |  | 2022/003/003/ |  |  |  |  |  |  |
|  |  |  |  | AMC1407 | in regulatory status. |  |  |  |
|  |  | 601/002 |  |  |  |  |  |  |
|  |  |  |  | AMC1508 | BACKGROUND: In 2020, MHRA |  |  |  |
|  |  |  |  | AMC1514 | granted LeMaitre permission |  |  |  |
|  |  |  |  | AMC1516 | (derogation) to supply AlboGraft |  |  |  |
|  |  |  |  | AMC1518 | devices on the UK market, without |  |  |  |
|  |  |  |  | AMC1520 | a CE mark. This derogation was |  |  |  |
|  |  |  |  | AMC1522 | requested because our previous |  |  |  |
|  |  |  |  | AMC1524 | notified body had stopped providing |  |  |  |
|  |  |  |  | AMC1608 | CE marking services and our new |  |  |  |
|  |  |  |  | AMC1609 | notified body had not completed |  |  |  |
|  |  |  |  | AMC1809 | their on-boarding review of the |  |  |  |
|  |  |  |  | AMC2010 | technical documentation. We are |  |  |  |
|  |  |  |  | AMC2011 | grateful to MHRA for allowing us |  |  |  |
|  |  |  |  | AMC2211 | to serve our customers while we |  |  |  |
|  |  |  |  | AMC2412 | worked to gain approval from our |  |  |  |
|  |  |  |  | AMC3006 | new notified body. MHRA has |  |  |  |
|  |  |  |  | AMC3008 | received a copy of this letter. |  |  |  |
|  |  |  |  | AMC3010 | LeMaitre has now received the CE |  |  |  |
|  |  |  |  | AMC3012 | mark approval from our new |  |  |  |
|  |  |  |  | AMC3014 | notified body for AlboGraft and we |  |  |  |
|  |  |  |  | AMC3016 | now have adequate capacity to |  |  |  |
|  |  |  |  | AMC3018 | serve all of our European customers |  |  |  |
|  |  |  |  | AMC3020 | with CE marked product. Although |  |  |  |
|  |  |  |  | AMC3022 | there is no safety risk with any of |  |  |  |
|  |  |  |  | AMC3024 | the derogated devices, we are |  |  |  |
|  |  |  |  | AMC4006 | required to withdraw any remaining |  |  |  |
|  |  |  |  | AMC4007 | devices as a condition of the |  |  |  |
|  |  |  |  | AMC4008 | derogation. |  |  |  |
|  |  |  |  | AMC6006 |  |  |  |  |
|  |  |  |  | AMC6007 |  |  |  |  |
|  |  |  |  | AMC6008 |  |  |  |  |
|  |  |  |  | AMC1520 |  |  |  |  |
|  |  |  |  | ASC3008 |  |  |  |  |
|  |  |  |  | ASC4006 |  |  |  |  |
|  |  |  |  | ASC4008 |  |  |  |  |
|  |  |  |  | ASC6006 |  |  |  |  |
|  |  |  |  | ASC6007 |  |  |  |  |
|  |  |  |  | ASC6008 |  |  |  |  |
| SOP 35-012 rev. B |  |  |  |  |  |  |  |  |

| Summary of Safety and Clinical Performance | Page 16 of 42 |  |
| --- | --- | --- |
| AlboGraft™ Polyester Vascular Graft |  |  |

| Date Initiated Recall # | Region | Products | Description |
| --- | --- | --- | --- |
| /FSCA |  | Involved |  |
| number |  |  |  |

ASC8006

ASC8007

ASC8008

ATC1407

ATC1518

ATC1522

ATC1608

ATC3012

ATC3014

ATC3016

ATC3020

Adverse Event Public Database Searches

There were 9 FDA MAUDE reports from 01 January 2019 to 31 May 2024 for the subject device known to the manufacturer or identified through a search of the FDA’s MAUDE database. The table below provides a summary of FDA MAUDE reports.

FDA MAUDE database - event type

| Event Types | # of Events | % of Events |
| --- | --- | --- |
| Death | 0 | 0% |
| Injury | 5 | 55.6% |
| Malfunction | 4 | 44.4% |
| Other or N/A | 0 | 0% |
| Total | 9 | 100% |

The most common device problem codes are summarized in the table below.

FDA MAUDE database – most frequent device problem codes

| Device Problem* | # of Events* | % of Events |
| --- | --- | --- |
| Material Puncture/Hole | 4 | 44.4% |
| Patient Device Interaction Problem | 2 | 22.2% |
| Device Contaminated at the User Facility | 2 | 22.2% |
| Material Rupture | 1 | 11.1% |
| Total | 9 | 100% |
| *Note: MAUDE allows more than one code per event. |  |  |

The most common patient problem codes are summarized in the table below.

SOP 35-012 rev. B

| Summary of Safety and Clinical Performance | Page 17 of 42 |  |
| --- | --- | --- |
| AlboGraft™ Polyester Vascular Graft |  |  |

FDA MAUDE database – most frequent patient problem codes

| Patient Problem* | # of Events* | % of Events |
| --- | --- | --- |
| No Clinical Signs, Symptoms or Conditions | 5 | 55.6% |
| Insufficient Information | 2 | 22.2% |
| Failure of Implant | 1 | 11.1% |
| No Consequences or Impact to Patient | 1 | 11.1% |
| Total | 9 | 100% |
| *Note: MAUDE allows more than one code per event. |  |  |

There were no reports of death. The table below provides additional details regarding the reports involving injury. Injuries were caused by the following device problems, material puncture/hole (n=2), patient device interaction problem (n=1), and material rupture (n=1).

FDA MAUDE database – reports involving injury

| Event | Device | Event Description / Manufacturer Narrative |
| --- | --- | --- |
| type | Problem |  |
|  | Code |  |
| Injury | Material | Event Description: During procedure, after completing distal anastomosis, |
|  | Puncture/Hole | surgeon observed a small hole in one of the bifurcated leg when flushing it with |
|  |  | saline. The section of the graft with the hole was discarded. There was no harm |
|  |  | to the patient as the result of this incident. |
|  |  | Manufacturer Narrative: We have not received the graft for evaluation since the |
|  |  | section of the graft containing the hole was discarded by the surgeon. The |
|  |  | remaining portion of the graft has been implanted in the patient. However, we |
|  |  | have observed the reported incident in the video that was provided to us. We |
|  |  | observed a small hole in the graft when user flushed the graft with saline. Our |
|  |  | review of the lot history records for this lot did not find any discrepancies either |
|  |  | in the manufacturing or packaging process that could be related to this incident. |
|  |  | During the manufacturing process, the exact graft was inspected by a quality |
|  |  | control inspector for holes, broken threads, and any loose fibers in the graft. No |
|  |  | defect was noted during the inspection process. A sample of grafts from this lot |
|  |  | number were also tested for water permeability and puncture test. All of the |
|  |  | samples met its specification. Further, we have not received any other complaints |
|  |  | of a similar nature from this lot. At this time, we remain inconclusive about the |
|  |  | root cause of the issue, but based on the documentation and complaint history |
|  |  | review, we do not believe there is a systemic issue with these grafts. It is possible |
|  |  | that this section of the graft came in contact with a sharp object during the |
|  |  | procedure which could have damaged the graft. Since January 2018- now, we |
|  |  | have received a total of 2 complaints where the surgeon observed a hole in the |
|  |  | graft. The current rate of occurrence of 0. 007% for this failure mode is within our |
|  |  | expected rate of occurrence of 0. 01%. None of the two reported cases caused any |
|  |  | harm to the patient. |

SOP 35-012 rev. B

| Summary of Safety and Clinical Performance | Page 18 of 42 |  |
| --- | --- | --- |
| AlboGraft™ Polyester Vascular Graft |  |  |

| Event | Device | Event Description / Manufacturer Narrative |  |
| --- | --- | --- | --- |
| type | Problem |  |  |
|  | Code |  |  |
| Injury | Patient Device | Event Description: Prosthesis thrombosis. Fem-pop arterial bypass. Need of |  |
|  | Interaction | another surgery to replace it. |  |
|  | Problem | Manufacturer Narrative: We could not perform a hands-on evaluation on the |  |
|  |  | graft since it has been discarded by the user facility. We have sent a list of follow- |  |
|  |  | up questions to the hospital but have not yet received a response despite our |  |
|  |  | repeated attempts. The investigation is ongoing, but our inquiries are not getting |  |
|  |  | any more response due to the increasing covid-19 situation in europe. We have |  |
|  |  | released a total of (b)(4) units of albograft vascular graft from this lot number. |  |
|  |  | This graft catalog# amc6006; lot# 211047 serial# (b)(4) was sold to the hospital |  |
|  |  | on (b)(6) 2018. We have not received any other complaints related to a similar |  |
|  |  | incident for this lot number. Our review of the lot history records for this graft did |  |
|  |  | not reveal any issues that could have contributed to this event. We have also |  |
|  |  | reviewed our complaint history for last 5 years. We did not find any other |  |
|  |  | complaints related to a similar issue reported to us by other hospitals. Please also |  |
|  |  | refer to manufacturer incident report# 1220948-2020-00106 relating to another |  |
|  |  | case of thrombosis that occurred at the same hospital after implanting albograft |  |
|  |  | vascular graft. We have received a total of (b)(4) complaints from this hospital on |  |
|  |  | the same day for this issue with the same exact description which raises questions |  |
|  |  | on how these grafts were prepared and implanted, including any postoperative |  |
|  |  | care provided to the patients. We have also reviewed our risk document. The |  |
|  |  | current rate of occurrence for this issue is within our expected rate. Our ifu |  |
|  |  | mentions a list of potential complications including thrombosis that occur with |  |
|  |  | the use of albograft vascular graft. |  |
| Injury | Material | Event Description: Spontaneous vertical rupture of the graft in the area of the |  |
|  | Rupture | proximal upper third in length, away from the anastomosis. |  |
|  |  | Manufacturer Narrative: We have received a piece of the graft that was |  |
|  |  | explanted by the surgeon. The piece was measured to be 5 cm in length. We |  |
|  |  | observed a semi-circular wear on the graft material at one end of the graft while |  |
|  |  | the other end appeared to be normal. The end where we observed the wear was |  |
|  |  | likely the section that presented the complications to the patient. However, it |  |
|  |  | could not be determined if the wear was due to the graft material deterioration or |  |
|  |  | if it occurred when surgeon explanted the graft. The graft was implanted in the |  |
|  |  | patient for more than 4 years. The graft did not show any signs of incorporation |  |
|  |  | to its surrounding tissue. Our review of the lot history records for this lot did not |  |
|  |  | find any discrepancies either in the manufacturing or packaging process that could |  |
|  |  | be related to this incident. During manufacturing process, the exact graft was |  |
|  |  | inspected by a quality control inspector as well as by a manufacturing staff for |  |
|  |  | holes, broken threads and any loose fibers in the graft. No defect was noted during |  |
|  |  | the inspection process. A sample of grafts from this lot number were also tested |  |
|  |  | for water permeability and puncture test. All of the samples met its specification. |  |
|  |  | Further, we have not received any other complaints of a similar nature from this |  |
|  |  | lot. Therefore, we believe that it was an isolated incident. During our follow-up |  |
|  |  | investigation, we learned: on (b)(6) 2020, the patient was presented with rest and |  |
|  |  | stress pains in the right thigh that have been present for about 10 days. The patient |  |
|  |  | was then admitted to the hospital by her general practitioner with leg vein |  |
|  |  | thrombosis. Patient pre-existing conditions include type ii diabetes mellitus, |  |

SOP 35-012 rev. B

| Summary of Safety and Clinical Performance | Page 19 of 42 |  |
| --- | --- | --- |
| AlboGraft™ Polyester Vascular Graft |  |  |

| Event | Device | Event Description / Manufacturer Narrative |
| --- | --- | --- |
| type | Problem |  |
|  | Code |  |
|  |  | hypertension, stress incontinence. Clinical findings on admission: externally no |
|  |  | hematoma discoloration on the thigh, pressure pain on the right groin, peripheral |
|  |  | pulses weak but palpable on both sides, polyneuropathy known, motor skills |
|  |  | unremarkable. Heart sounds purely rhythmic, vesicular breathing sound on both |
|  |  | sides, abdominal wall soft no pressure pain, no defense tension. Sonography: |
|  |  | arterial false aneurysm right groin resting ecg: sinus rhythm, left type, heart rate |
|  |  | 83 / min, rs envelope v3 / 4, no significant regression disorders. Ct pelvic leg |
|  |  | angiography with contrast was performed on (b)(6) 2020. The visceral arteries |
|  |  | were filling with contract. Moderate to high grade stenosis on the celiac trunk with |
|  |  | post-stenotic dilation, possibly functional. Ams and renal arteries were properly |
|  |  | displayed during angiography. On the right thigh which presented the |
|  |  | complication, surgeon diagnosed an occlusion of the fem-pop bypass, a long- |
|  |  | stretched liquid cavity is shown with different density values and contrast medium |
|  |  | absorption at the edge. Multiple air pockets in the area after surgery. In the distal |
|  |  | portion fluid collection in the sartorius muscle, starting from the bypass height, |
|  |  | repeatedly septated. Here also contrast agent uptake on the margin. The distal |
|  |  | anastomosis is presumably refilled retrograde. Contrast infusion of the artery |
|  |  | profunda femoris. Narrow-caliber popliteal artery with semicircular calcifications |
|  |  | and moderate stenosis. Exit of a filiform anterior tibial artery, which only appears |
|  |  | to be filled with contrast proximally. Significant calcifications of the tibiofibular |
|  |  | trunk. Division into a narrow posterior tibial artery and fibular artery. Restricted |
|  |  | two-vessel supply. Inguinal drainage over the right, which ends laterally of the |
|  |  | fluid cavity. On the left thigh, the surgeon also noticed a semi-circular |
|  |  | calcification and installation of a fem-pop bypass. Circular compression seam |
|  |  | around the entire bypass. The bypass displayed well in itself. The distal portion |
|  |  | shows fluid collection from the bypass, holding contrast on the edge towards |
|  |  | mediodorsal (density values over 20 he). Maximum axial expansion 2. 7 cm. |
|  |  | Contrast into the artery profunda femoris. Popliteal artery narrow-caliber with |
|  |  | moderate calcifications. Branch of the anterior tibial artery, also only filiform. |
|  |  | Truncus tibio-fibularis with multiple calcifications. Narrow posterior tibial artery |
|  |  | and fibular artery. Severely restricted two-vessel supply. Evaluation: the cause |
|  |  | was determined to be occlusion of the fem-pop bypass graft (albograft vascular |
|  |  | graft) on the right thigh. Periprosthetic, long-distance fluid collection (image- |
|  |  | morphological suspected abscess), after consultation with the colleague of the |
|  |  | vascular surgery department, there were multiple times occurred hematomas / |
|  |  | seromas, also in the sartorius muscle. Severely restricted two-vessel supply to the |
|  |  | right lower leg. Perfused fem. -pop. Bypass on the left. Periprosthetic compression |
|  |  | ligature. At the level of the distal anastomosis, fluid collection is evident, as well |
|  |  | as on the opposite side. Restricted two-vessel supply to the left lower leg. The |
|  |  | section of the ruptured albograft was then explanted on (b)(6) 2020. The operation |
|  |  | was carried under general anesthesia. Intraoperatively, the damaged piece of the |
|  |  | femoro-popliteal bypass graft was removed and the ends were then ligated. The |
|  |  | patient's symptoms improved significantly immediately after the surgery. The |
|  |  | wound was always irritable and showed a primary tendency to heal. The |
|  |  | peripheral motor skills and sensitivity were intact at all times. An angio-ct showed |
|  |  | the finding of a periprosthetic long-distance fluid retention on the right |
|  |  | corresponding to the intraoperative finding of a partly old, partly fresh hematoma. |
|  |  | At the time of discharge, the patient was independently mobile at the ward level |

SOP 35-012 rev. B

| Summary of Safety and Clinical Performance | Page 20 of 42 |  |
| --- | --- | --- |
| AlboGraft™ Polyester Vascular Graft |  |  |

| Event | Device | Event Description / Manufacturer Narrative |
| --- | --- | --- |
| type | Problem |  |
|  | Code |  |
|  |  | and also free of symptoms. On (b)(6) 2020, patient was discharged inpatient |
|  |  | treatment without symptoms. Patient was recommended: regular wound control |
|  |  | and dressing changes. Suture removal after 14 days. Pain-adapted analgesia. |
|  |  | Check-up with the physician in 2 weeks for clinical control and planning of further |
|  |  | therapy. |
| Injury | Material | Event Description: Intraoperatively, a hole was noticed in the graft. The hole |
|  | Puncture/Hole | was 3-4mm in size. On one picture is a 1. 2mm dull- headed cannula in the hole |
|  |  | of the graft. The hole was sewn over. The patient was still in the operation room, |
|  |  | when they called us. There was no impact on patient's health as the result of this |
|  |  | incident. |
|  |  | Manufacturer Narrative: We have not received the graft for evaluation since the |
|  |  | graft was implanted in the patient. However, we have observed the reported |
|  |  | incident in the pictures provided to us. We observed a small hole in the graft. The |
|  |  | surgeon did not check for holes or other defects before implantation. Our review |
|  |  | of the lot history records for this lot did not find any discrepancies either in the |
|  |  | manufacturing or packaging process that could be related to this incident. During |
|  |  | manufacturing process, the exact graft was inspected by a quality control |
|  |  | inspector for holes, broken threads and any loose fibers in the graft. No defect was |
|  |  | noted during the inspection process. A sample of grafts from this lot number were |
|  |  | also tested for water permeability and puncture test. All of the samples met its |
|  |  | specification. Further, we have not received any other complaints of a similar |
|  |  | nature from this lot. Therefore, we believe that it was an isolated incident. We |
|  |  | have also reviewed our complaint history records for last 4 years and searched for |
|  |  | a similar issue where a hole/tear was found in an albograft polyester vascular graft. |
|  |  | We have not received any other complaints for a similar issue. The current rate of |
|  |  | occurrence is within our expected rate of occurrence. Unfortunately, our analysis |
|  |  | was limited to only the information provided and no physical sample was |
|  |  | available to us for investigation. We therefore remain inconclusive about the root |
|  |  | cause of the issue, but based on the documentation and complaint history review, |
|  |  | we do not believe there is a systemic issue with these devices. No corrective action |
|  |  | is needed at this time. Trends are monitored on a monthly basis and if action is |
|  |  | required, appropriate investigation will be performed. Surgeon sewed the hole |
|  |  | after observing the hole in the graft. Approximately 150 cc of blood was lost and |
|  |  | the operation time extended by 20 minutes. On may 19, 2020, we were informed |
|  |  | that the patient is doing fine. There were no impact on the patient's health as a |
|  |  | result of the graft malfunction. |

- Summary of clinical evaluation and post-market clinical follow-up (PMCF)

  - Summary of clinical data related to equivalent device, if applicable: No equivalency used in the assessment of these devices.

  - Summary of clinical data from conducted investigations of the device

before the CE-marking, if applicable

There were no manufacturer-sponsored pre-market investigations conducted with the device. The AlboGraft Polyester Vascular Patch was previously manufactured

SOP 35-012 rev. B

| Summary of Safety and Clinical Performance | Page 21 of 42 |  |
| --- | --- | --- |
| AlboGraft™ Polyester Vascular Graft |  |  |

by Biomateriali S.r.l., a subsidiary of LeMaitre Vascular Inc., in Brindisi, Italy. The device was first approved for CE mark under LeMaitre Vascular Inc. in 2010. The Post-Market Clinical Follow-up studies performed for the subject device are listed below:

- Biomateriali Albograft™- A Retrospective Clinical Data Review (2008)

- Biomateriali Albograft™ Thoracic Aortic Application A Retrospective Clinical Data Review (2009)

- Biomateriali AlboGraft™ - A Retrospective Clinical Data Review (2010)

The findings of these studies are summarized below.

The product is a mature product currently on the market for a well-established intended use. It has been developed by incremental changes and is based on the Hemashield Microvel Double Velour Knitted and Woven Vascular Graft. The AlboGraft was previously manufactured by Biomateriali S. r. l., a subsidiary of LeMaitre Vascular, Inc. in Brindisi, Itay. LeMaitre Vascular has transferred the manufacturing from Italy to Burlington, MA.

Biomaterali conducted the following three studies:

- A Retrospective Clinical Data Review (2008), which concluded that “overall, aortic reconstruction with the Albograft performed favourably compared with existing literature. Albograft has comparable short and long term patency, and also exhibits similar mortality and morbity rates in the follow-up period, when compared with existing literature. We did not observe any adverse events which were directly related to the Albograft Polyester prosthesis. The overall handling in terms of suturing, conformability to the anastomosis, and suture hole bleeding compares well with polyester grafts of other brands previously used (C.R.BARD Dialine II; Boston Scientific Hemashield). One of the major reasons we originally switched to the Albograft was a competitive price offer by the current distributor.”

- Thoracic Aortic Application A Retrospective Clinical Data Review (2009), which concluded “Overall, aortic reconstruction with the Albograft performed favourably compared with existing literature. Albograft has comparable short and long term patency and also exhibits similar mortality and morbidly rates in the follow-up period when compared with existing literature. We did not observe any adverse events which were directly related to the AlboGraft Polyester prosthesis.”

- A Retrospective Clinical Data Review (2010), which concluded that “Overall, the peripheral intraoperative reconstructions and outcomes with the AlboGraft described in this report performed well when compared with existing literature. The AlboGraft had comparable immediate (30 day) and long term (24 month) patency and also exhibits similar mortality and morbidity rates in the follow-up period up to max. 36 month when compared with existing literature. We did not observe any

SOP 35-012 rev. B

| Summary of Safety and Clinical Performance | Page 22 of 42 |  |
| --- | --- | --- |
| AlboGraft™ Polyester Vascular Graft |  |  |

adverse events which were directly related to the AlboGraft Polyester prosthesis.”

iii)	Summary of clinical data from other sources, if applicable

A complete summary of performance and clinical benefit outcomes for the device under evaluation in comparison to the state-of-the-art acceptance criteria is presented in the table below. At up to 2-years follow-up, primary patency (88.7-94.8%), secondary patency (91.9-95.1%), survival (89.6-99.0%), limb salvage (98.4%), and reintervention rates (reported as reoperation for bleeding; 2.9%) for the subject devices were comparable to similar treatments and within the acceptable limits reported in the state-of-the-art literature evaluation for the treatment of aneurysmal and occlusive disease. There were two instances when the acceptance criterion was not met. Following the treatment of occlusive disease of the iliac arteries in post-market study C, the survival rate at 30 days (62/66, 93.9%) did not meet the acceptance criterion (≥ 96.5%). However, within the same study the 1-year survival rate (59/61, 96.7%) met the acceptance criterion (≥ 83.9%). Of note, this is due to one patient being lost to follow-up (i.e., n=62 patients were discharged from hospital alive after 30 days and n=1 patient was discharged to a psychiatric hospital with no more FU visits. The acceptance criteria for other measures, primary patency and limb salvage, were met in the same study. Following the treatment of aneurysmal or occlusive disease in the retrospective PMCF study, the freedom from reintervention rate at 1 year (88.7%), did not meet the acceptance criterion established from the inverse of the reintervention benchmark (91.7%, 95% CI: 89.5-93.7%). However, the rate is similar to the lower end of the 95% confidence interval (89.5%).

Comparison of performance outcomes to acceptance criteria

| Outcome | Indication | Location | Device | Follow-up | Benchmark | Conclusions |
| --- | --- | --- | --- | --- | --- | --- |
|  |  |  | under |  |  |  |
|  |  |  | evaluation |  |  |  |
| Survival | Aneurysmal | Infrarenal | 99.0% | Perioperative | ≥ 89.1% | Acceptance |
|  | disease | artery | (101/102)* |  |  | criteria met |
|  |  |  | 98.0% | 30 days | ≥ 89.1% | Acceptance |
|  |  |  | (100/102)* |  |  | criteria met |
|  |  |  | 96.9% | 1 year | ≥ 81.6% | Acceptance |
|  |  |  | (93/96)* |  |  | criteria met |
|  |  |  | 94.8% | 2 years | ≥ 81.6% | Acceptance |
|  |  |  | (91/96)* |  |  | criteria met |
|  | Aneurysmal | Ascending | 94.8% | 30 days | ≥ 89.1% | Acceptance |
|  | disease** | or | (73/77)* |  |  | criteria met |
|  |  | descending | 89.6% | 1 year | ≥ 81.6% | Acceptance |
|  |  | aorta | (69/77)* |  |  | criteria met |
|  | Occlusive | Iliac artery | 93.9% | 30 days | ≥ 96.5% | Acceptance |
|  | disease*** |  | (62/66)* |  |  | criteria not met |
|  |  |  | 96.7% | 1 year | ≥ 83.9% | Acceptance |
|  |  |  | (59/61)* |  |  | criteria met |
|  | Aneurysmal | N/A | 91.9% | 1 year | ≥ 83.9% | Acceptance |
|  | or occlusive |  | (203/221) |  |  | criteria met |
|  | disease |  |  |  |  |  |
| Primary patency | Aneurysmal | Infrarenal | 90.0% | At hospital | ≥ 87.9% | Acceptance |
|  | disease | artery | (90/100) | discharge |  | criteria met |

SOP 35-012 rev. B

| Summary of Safety and Clinical Performance | Page 23 of 42 |  |
| --- | --- | --- |
| AlboGraft™ Polyester Vascular Graft |  |  |

| Outcome | Indication | Location | Device | Follow-up | Benchmark | Conclusions |
| --- | --- | --- | --- | --- | --- | --- |
|  |  |  | under |  |  |  |
|  |  |  | evaluation |  |  |  |
|  |  |  | 94.8% | 1 year | ≥ 87.9% | Acceptance |
|  |  |  | (91/96) |  |  | criteria met |
|  | Occlusive | Iliac artery | 88.9% | 1 year | ≥ 77.0% | Acceptance |
|  | disease*** |  |  |  |  | criteria met |
|  |  |  | 88.9% | 2 years | ≥ 77.0% | Acceptance |
|  |  |  |  |  |  | criteria met |
|  | Aneurysmal | N/A | 88.7% | 1 year | ≥ 77.0% | Acceptance |
|  | or occlusive |  | (196/221) |  |  | criteria met |
|  | disease |  |  |  |  |  |
| Secondary patency | Occlusive | Iliac artery | 95.1% | 1 year | ≥ 77.8% | Acceptance |
|  | disease*** |  |  |  |  | criteria met |
|  |  |  | 95.1% | 2 years | ≥ 77.8% | Acceptance |
|  |  |  |  |  |  | criteria met |
|  | Aneurysmal | N/A | 91.9% | 1 year | ≥ 77.8% | Acceptance |
|  | or occlusive |  | (203/221) |  |  | criteria met |
|  | disease |  |  |  |  |  |
| Limb salvage | Occlusive | Iliac artery | 98.4% | 1 year | ≥ 79.4%† | Acceptance |
|  | disease*** |  |  |  |  | criteria met |
| Reoperation for | Aneurysmal | Infrarenal | 2.9% | Perioperative | ≤ 12.4% | Acceptance |
| bleeding | disease | artery | (3/102) |  |  | criteria met |
| Freedom from | Aneurysmal | N/A | 88.7% | 1 year | ≥ 91.7‡ | Acceptance |
| reintervention | or occlusive |  | (196/221) |  |  | criteria not met |
|  | disease |  |  |  |  |  |
| Technical success | Aneurysmal | N/A | 100% | Not reported | ≥ 98.7% | Acceptance |
|  | or occlusive |  | (221/221) |  |  | criteria met |
|  | disease |  |  |  |  |  |
| Freedom from | Aneurysmal | N/A | 76% | Not reported | N/A | Acceptance |
| procedure-related | or occlusive |  | (169/221) |  |  | criteria met |
| (severe) adverse | disease |  |  |  |  |  |
| events |  |  |  |  |  |  |
| Freedom from | Aneurysmal | N/A | 92% | 30 days | N/A | Acceptance |
| device-related | or occlusive |  | (203/221) |  |  | criteria met |
| (severe) adverse | disease | N/A | 86% | 30 days – 1 | N/A | Acceptance |
| events |  |  | (189/221) | year |  | criteria met |

* Survival was calculated from reported mortality rate

† Acceptance criterion for this outcome was based on reported amputation rates in the state of the art.

‡ Acceptance criterion for this outcome was calculated as the inverse of reintervention rates in the state of the art ** 99% of patients were treated for thoracic aortic aneurysms or acute or chronic dissections of the aorta.

*** 97% of patients were treated for occlusive disease.

The table below provides direct comparisons between the safety outcomes for the device under evaluation (from post-market studies; and the acceptance criteria for those outcomes established by the state of the art assessment. None of the adverse events reported in the 3 post-market studies were directly attributed to the

SOP 35-012 rev. B

| Summary of Safety and Clinical Performance | Page 24 of 42 |  |
| --- | --- | --- |
| AlboGraft™ Polyester Vascular Graft |  |  |

device under evaluation. However, the PMCF retrospective registry observed a total of 21 device-related complications in 18 patients throughout the 30-day follow-up period. The two most common device-related adverse events besides lymphocele (8/221, 4%) were graft infection (4/221, 2%) and graft thrombosis (4/221, 2%) requiring intervention. Graft infection (3%), graft thrombosis (3%) and superficial wound infections (3%) were the most frequent complication during follow-up between 30-days and 1 year. In this period, one death occurred due to graft infection. In addition, many of the adverse events were not reported in the clinical studies from the state of the art, and therefore there were no acceptance criteria established for those outcomes. Per the risk files, these risks are reduced as much as possible. Mortality, specific rates of myocardial infarction (e.g., cardiac complications) thrombosis, and amputation rates met the acceptance criteria established by the state of the art. No graft integrity outcomes (e.g., graft dilation, graft disintegration, etc.) were reported in the post-market studies or PMCF registry. There were three outcomes that did not meet the acceptance criteria as established by the state of the art. In the treatment of occlusive disease of the iliac arteries in post-market Study C, the 30-day mortality rate (6.1%), but not the 1 year mortality rate (3.5%), exceeded the acceptance criteria. None of the deaths were attributed to the graft implanted. In the treatment of aneurysmal disease, the acceptance criteria were not met for visceral ischemia (post-market Study B, 3.9%). In addition, the acceptance criteria were not met for wound hematoma in a separate study (post-market Study C, 6.1%). In both instances, visceral ischemia and wound hematoma, the outcomes were not attributed to the graft implanted. In addition, the symptoms for wound hematoma were either treated conservatively or successfully via revision surgery with no further sequelae. One of the 3 patients experiencing visceral ischemia did experience multi-organ failure and death and the other 2 patients recovered without complications. The adverse events from these studies were reviewed via the risk management process of the manufacturer to determine whether the list of potential safety outcomes associated with the device under evaluation in the IFU is complete.

Comparison of safety outcomes to acceptance criteria

| Residual | Indication | Location | Device under |  | Device | Follow-up | Benchmark | Conclusions |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| risk in RMF |  |  | evaluation |  | under |  |  |  |
| and IFU |  |  | (literature, |  | evaluation |  |  |  |
|  |  |  | investigations, |  | (complaints/ |  |  |  |
|  |  |  | PMCF, |  | vigilance) |  |  |  |
|  |  |  | registries) |  |  |  |  |  |
| Gluteal | Aneurysmal | Infrarenal | 1.0% | (1/102) | NR | Perioperative | NR | N/A |
| necrosis | disease | artery |  |  |  |  |  |  |
| Mortality | Aneurysmal | Infrarenal | 1.0% | (1/102) | NR | Perioperative | ≤ 3.8% | Acceptance |
|  | disease | artery |  |  |  |  |  | criteria met |
|  |  |  | 2.0% | (2/102) | NR | 30 days | ≤ 6.8% | Acceptance |
|  |  |  |  |  |  |  |  | criteria met |
|  |  |  | 3.3% | (3/96) | NR | 1 year | ≤ 15.5% | Acceptance |
|  |  |  |  |  |  |  |  | criteria met |
|  |  |  | 5.2% | (5/96) | NR | 2 years | ≤ 15.5% | Acceptance |
|  |  |  |  |  |  |  |  | criteria met |
|  | Aneurysmal | Ascending | 5.2% | (4/77) | NR | 30 days | ≤ 6.8% | Acceptance |
|  | disease** | or |  |  |  |  |  | criteria met |
|  |  | descending | 10.4% (8/77) |  | NR | 1 year | ≤ 15.5% | Acceptance |
|  |  | aorta |  |  |  |  |  | criteria met |

SOP 35-012 rev. B

| Summary of Safety and Clinical Performance | Page 25 of 42 |  |
| --- | --- | --- |
| AlboGraft™ Polyester Vascular Graft |  |  |

| Residual | Indication | Location | Device under | Device | Follow-up | Benchmark | Conclusions |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| risk in RMF |  |  | evaluation | under |  |  |  |  |
| and IFU |  |  | (literature, | evaluation |  |  |  |  |
|  |  |  | investigations, | (complaints/ |  |  |  |  |
|  |  |  | PMCF, | vigilance) |  |  |  |  |
|  |  |  | registries) |  |  |  |  |  |
|  | Occlusive | Iliac artery | 6.1% (4/66) | NR | 30 days | ≤ 3.5% | Acceptance |  |
|  | disease*** |  |  |  |  |  | criteria not |  |
|  |  |  |  |  |  |  | met |  |
|  |  |  | 3.3% (2/61) | NR | 1 year | ≤ 16.1% | Acceptance |  |
|  |  |  |  |  |  |  | criteria met |  |
|  | Aneurysmal | N/A | 3% (7/221) | NR | 30 days | ≤ 3.5% | Acceptance |  |
|  | or occlusive |  |  |  |  |  | criteria met |  |
|  | disease | N/A | 8.1% (18/221) | NR | 1 year | ≤ 15.5% | Acceptance |  |
|  |  |  |  |  |  |  | criteria met |  |
| Device- | Aneurysmal | N/A | 0% (0/221) | NR | 30 days | NR | N/A |  |
| related | or occlusive | N/A | 0.45% (1/221) | NR | 1 year | NR | N/A |  |
| mortality | disease |  |  |  |  |  |  |  |
| Postoperative | Aneurysmal | N/A | 23.5% | NR | 30 days | NR | N/A |  |
| procedure- | or occlusive |  | (52/221) |  |  |  |  |  |
| related | disease |  |  |  |  |  |  |  |
| complication† |  |  |  |  |  |  |  |  |
| Postoperative | Aneurysmal | N/A | 8% (18/221) | NR | 30 days | NR | N/A |  |
| device- | or occlusive |  |  |  |  |  |  |  |
| related | disease |  |  |  |  |  |  |  |
| complication‡ |  |  |  |  |  |  |  |  |
| Ileus | Aneurysmal | Infrarenal | 2.9% (3/102) | NR | Perioperative | NR | N/A |  |
|  | disease | artery |  |  |  |  |  |  |
| Ileus† | Aneurysmal | N/A | 1% (2/221) | NR | 30 days | NR | N/A |  |
|  | or occlusive |  |  |  |  |  |  |  |
|  | disease |  |  |  |  |  |  |  |
| Myocardial | Aneurysmal | Infrarenal | 1% (1/102) | NR | Perioperative | ≤ 5.2% | Acceptance |  |
| infarction | disease | artery |  |  |  |  | criteria met |  |
|  | Aneurysmal | Ascending | 2.7% (2/73) | NR | 1 year | ≤ 5.2% | Acceptance |  |
|  | disease** | or |  |  |  |  | criteria met |  |
|  |  | descending |  |  |  |  |  |  |
|  |  | aorta |  |  |  |  |  |  |
| Myocardial | Aneurysmal | N/A | 1% (3/221) | NR | 30 days | ≤ 5.2% | Acceptance |  |
| infarction† | or occlusive |  |  |  |  |  | criteria met |  |
|  | disease |  |  |  |  |  |  |  |
| Myocardial |  | N/A | 0% (0/221) | NR | 30 days | ≤ 5.2% | Acceptance |  |
| infarction‡ |  |  |  |  |  |  | criteria met |  |
|  |  | N/A | 0% (0/221) | NR | 30 days – 1 | ≤ 5.2% | Acceptance |  |
|  |  |  |  |  | year |  | criteria met |  |

SOP 35-012 rev. B

| Summary of Safety and Clinical Performance | Page 26 of 42 |  |
| --- | --- | --- |
| AlboGraft™ Polyester Vascular Graft |  |  |

| Residual | Indication | Location | Device under | Device | Follow-up | Benchmark | Conclusions |
| --- | --- | --- | --- | --- | --- | --- | --- |
| risk in RMF |  |  | evaluation | under |  |  |  |
| and IFU |  |  | (literature, | evaluation |  |  |  |
|  |  |  | investigations, | (complaints/ |  |  |  |
|  |  |  | PMCF, | vigilance) |  |  |  |
|  |  |  | registries) |  |  |  |  |
| Bleeding | Aneurysmal | Infrarenal | 2.0% (2/102) | NR | Perioperative | NR | N/A |
|  | disease | artery |  |  |  |  |  |
|  | Aneurysmal | Ascending | 2.6% (2/77) | NR | Perioperative | NR | N/A |
|  | disease** | or |  |  |  |  |  |
|  |  | descending |  |  |  |  |  |
|  |  | aorta |  |  |  |  |  |
| Bleeding† | Aneurysmal | N/A | 3% (7/221) | NR | 30 days | NR | N/A |
|  | or occlusive |  |  |  |  |  |  |
|  | disease |  |  |  |  |  |  |
| Paraparesis | Aneurysmal | Infrarenal | 1% (1/102) | NR | Perioperative | NR | N/A |
|  | disease | artery |  |  |  |  |  |
|  | Aneurysmal | Ascending | 1.4% (1/73) | NR | 1 year | NR | N/A |
|  | disease** | or |  |  |  |  |  |
|  |  | descending |  |  |  |  |  |
|  |  | aorta |  |  |  |  |  |
| Wound | Aneurysmal | Infrarenal | 1% (1/102) | NR | Perioperative | NR | N/A |
| infection | disease | artery |  |  |  |  |  |
|  | Aneurysmal | Ascending | 3.9% (3/77) | NR | Perioperative | NR | N/A |
|  | disease** | or |  |  |  |  |  |
|  |  | descending |  |  |  |  |  |
|  |  | aorta |  |  |  |  |  |
| Superficial | Aneurysmal | N/A | 1% (3/221) | NR | 30 days | NR | N/A |
| wound | or occlusive | N/A | 3% (7/221) | NR | 30 days – 1 | NR | N/A |
| infection‡ | disease |  |  |  | year |  |  |
| Acute kidney | Aneurysmal | Ascending | 1.3% (1/77) | NR | Perioperative | ≤ 9.1% | Acceptance |
| failure | disease** | or |  |  |  |  | criteria met |
|  |  | descending |  |  |  |  |  |
|  |  | aorta |  |  |  |  |  |
| Kidney | Aneurysmal | N/A | 9% (19/221) | NR | 30 days | ≤ 9.1% | Acceptance |
| failure† | or occlusive |  |  |  |  |  | criteria met |
|  | disease |  |  |  |  |  |  |
| MOF† | Aneurysmal | N/A | 3% (6/221) | NR | 30 days | NR | N/A |
|  | or occlusive |  |  |  |  |  |  |
|  | disease |  |  |  |  |  |  |
| Chylothorax | Aneurysmal | Ascending | 2.6% (2/77) | NR | Perioperative | NR | N/A |
|  | disease** | or |  |  |  |  |  |
|  |  | descending |  |  |  |  |  |
|  |  | aorta |  |  |  |  |  |

SOP 35-012 rev. B

| Summary of Safety and Clinical Performance | Page 27 of 42 |  |
| --- | --- | --- |
| AlboGraft™ Polyester Vascular Graft |  |  |

| Residual | Indication | Location | Device under |  | Device | Follow-up | Benchmark | Conclusions |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| risk in RMF |  |  | evaluation |  | under |  |  |  |
| and IFU |  |  | (literature, |  | evaluation |  |  |  |
|  |  |  | investigations, |  | (complaints/ |  |  |  |
|  |  |  | PMCF, |  | vigilance) |  |  |  |
|  |  |  | registries) |  |  |  |  |  |
| Compartment | Aneurysmal | Ascending | 2.6% | (2/77) | NR | Perioperative | NR | N/A |
| syndrome | disease** | or |  |  |  |  |  |  |
|  |  | descending |  |  |  |  |  |  |
|  |  | aorta |  |  |  |  |  |  |
| Hemiparesis | Aneurysmal | Ascending | 3.9% | (3/77) | NR | Perioperative | NR | N/A |
|  | disease** | or |  |  |  |  |  |  |
|  |  | descending |  |  |  |  |  |  |
|  |  | aorta |  |  |  |  |  |  |
| Paravalvular | Aneurysmal | Ascending | 2.6% | (2/77) | NR | Perioperative | NR | N/A |
| leak | disease** | or |  |  |  |  |  |  |
|  |  | descending |  |  |  |  |  |  |
|  |  | aorta |  |  |  |  |  |  |
| Transitory | Aneurysmal | Ascending | 3.9% | (3/77) | NR | Perioperative | NR | N/A |
| psychotic | disease** | or |  |  |  |  |  |  |
| syndrome |  | descending |  |  |  |  |  |  |
|  |  | aorta |  |  |  |  |  |  |
| Visceral | Aneurysmal | Ascending | 3.9% | (3/77) | NR | Perioperative | ≤ 2.4% | Acceptance |
| ischemia | disease** | or |  |  |  |  |  | criteria not |
|  |  | descending |  |  |  |  |  | met |
|  |  | aorta |  |  |  |  |  |  |
| Mesenteric | Aneurysmal | N/A | 0.5% | (1/221) | NR | 30 days | ≤ 2.4% | Acceptance |
| ischemia† | or occlusive |  |  |  |  |  |  | criteria met |
|  | disease |  |  |  |  |  |  |  |
| Local | Aneurysmal | Ascending | 2.6% | (2/77) | NR | Perioperative | NR | N/A |
| nonvascular | disease** | or |  |  |  |  |  |  |
| complications |  | descending |  |  |  |  |  |  |
| (not |  | aorta |  |  |  |  |  |  |
| contribute to | Occlusive | Iliac artery | 8.2% | (5/61) | NR | 1 year | NR | N/A |
| morbidity) | disease*** |  |  |  |  |  |  |  |
| Aneurysm | Aneurysmal | Ascending | 1.4% | (1/73) | NR | 1 year | NR | N/A |
| growth | disease** | or |  |  |  |  |  |  |
|  |  | descending |  |  |  |  |  |  |
|  |  | aorta |  |  |  |  |  |  |
| Atrial | Aneurysmal | Ascending | 2.7% | (2/73) | NR | 1 year | ≤ 4.7% | Acceptance |
| fibrillation | disease** | or |  |  |  |  |  | criteria met |
|  |  | descending |  |  |  |  |  |  |
|  |  | aorta |  |  |  |  |  |  |
| Descending | Aneurysmal | Ascending | 1.4% | (1/73) | NR | 1 year | NR | N/A |
| aorta | disease** | or |  |  |  |  |  |  |
| dissection |  | descending |  |  |  |  |  |  |
|  |  | aorta |  |  |  |  |  |  |

SOP 35-012 rev. B

| Summary of Safety and Clinical Performance | Page 28 of 42 |  |
| --- | --- | --- |
| AlboGraft™ Polyester Vascular Graft |  |  |

| Residual | Indication | Location | Device under | Device | Follow-up | Benchmark | Conclusions |
| --- | --- | --- | --- | --- | --- | --- | --- |
| risk in RMF |  |  | evaluation | under |  |  |  |
| and IFU |  |  | (literature, | evaluation |  |  |  |
|  |  |  | investigations, | (complaints/ |  |  |  |
|  |  |  | PMCF, | vigilance) |  |  |  |
|  |  |  | registries) |  |  |  |  |
| Deep vein | Aneurysmal | Ascending | 1.4% (1/73) | NR | 1 year | ≤ 6.6% | Acceptance |
| thrombosis | disease** | or |  |  |  |  | criteria met |
|  |  | descending |  |  |  |  |  |
|  |  | aorta |  |  |  |  |  |
| Sternum | Aneurysmal | Ascending | 1.4% (1/73) | NR | 1 year | NR | N/A |
| instability | disease** | or |  |  |  |  |  |
|  |  | descending |  |  |  |  |  |
|  |  | aorta |  |  |  |  |  |
| Bypass | Occlusive | Iliac artery | 6.1% (4/66) | NR | Perioperative | NR | N/A |
| occlusions | disease*** |  | 3.3% (2/61) | NR | 1 year | NR | N/A |
| Wound | Occlusive | Iliac artery | 6.1% (4/66) | NR | Perioperative | ≤ 2.4% | Acceptance |
| hematoma | disease*** |  |  |  |  |  | criteria not |
|  |  |  |  |  |  |  | met |
| Hematoma‡ | Aneurysmal | N/A | 1% (2/221) | NR | 30 days | ≤ 2.4% | Acceptance |
|  | or occlusive |  |  |  |  |  | criteria met |
|  | disease |  |  |  |  |  |  |
| Groin | Occlusive | Iliac artery | 1.5% (1/66) | NR | Perioperative | NR | N/A |
| Seroma | disease*** |  |  |  |  |  |  |
| Urinary tract | Occlusive | Iliac artery | 1.5% (1/66) | NR | Perioperative | NR | N/A |
| infection | disease*** |  |  |  |  |  |  |
| Amputation | Occlusive | Iliac artery | 1.6% (1/61) | NR | 1 year | ≤ 20.6% | Acceptance |
|  | disease*** |  |  |  |  |  | criteria met |
| Loss of limb† | Aneurysmal | N/A | 1% (2/221) | NR | 30 days | ≤ 8.1% | Acceptance |
|  | or occlusive |  |  |  |  |  | criteria met |
| Loss of limb‡ | disease | N/A | 0% (0/221) | NR | 30 days – 1 | ≤ 20.6% | Acceptance |
|  |  |  |  |  | year |  | criteria met |
| Angina | Occlusive | Iliac artery | 1.6% (1/61) | NR | 1 year | NR | N/A |
| abdominalis | disease*** |  |  |  |  |  |  |
| Infection | Occlusive | Iliac artery | 0% (0/61) | NR | 1 year | NR | N/A |
|  | disease*** |  |  |  |  |  |  |
| Graft | Aneurysmal | N/A | 2% (4/221) | NR | 30 days | NR | N/A |
| infection‡ | or occlusive | N/A | 3% (7/221) | NR | 30 days – 1 | NR | N/A |
|  | disease |  |  |  | year |  |  |
| Graft | Aneurysmal | N/A | 3% (6/168) | 0.004% | 1 year | NR | N/A |
| thrombosis | or occlusive |  |  | (2/47,011) |  |  |  |
|  | disease |  |  |  |  |  |  |
| Thrombosis‡ | Aneurysmal | N/A | 2% (4/221) | NR | 30 days | ≤ 6.6% | Acceptance |
|  | or occlusive |  |  |  |  |  | criteria met |
|  | disease | N/A | 3% (6/221) | NR | 30 days – 1 | ≤ 6.6% | Acceptance |
|  |  |  |  |  | year |  | criteria met |

SOP 35-012 rev. B

| Summary of Safety and Clinical Performance | Page 29 of 42 |  |
| --- | --- | --- |
| AlboGraft™ Polyester Vascular Graft |  |  |

| Residual | Indication | Location | Device under | Device | Follow-up | Benchmark | Conclusions |
| --- | --- | --- | --- | --- | --- | --- | --- |
| risk in RMF |  |  | evaluation | under |  |  |  |
| and IFU |  |  | (literature, | evaluation |  |  |  |
|  |  |  | investigations, | (complaints/ |  |  |  |
|  |  |  | PMCF, | vigilance) |  |  |  |
|  |  |  | registries) |  |  |  |  |
| ARDS† | Aneurysmal | N/A | 1% (2/221) | NR | 30 days | NR | N/A |
|  | or occlusive |  |  |  |  |  |  |
|  | disease |  |  |  |  |  |  |
| Heart failure† | Aneurysmal | N/A | 1% (2/221) | NR | 30 days | ≤ 4.7% | Acceptance |
|  | or occlusive |  |  |  |  |  | criteria met |
|  | disease |  |  |  |  |  |  |
| Nerve injury† | Aneurysmal | N/A | 0% (0/221) | NR | 30 days | NR | N/A |
|  | or occlusive |  |  |  |  |  |  |
|  | disease |  |  |  |  |  |  |
| Paraplegia† | Aneurysmal | N/A | 0.5% (1/221) | NR | 30 days | ≤ 3.7% | Acceptance |
|  | or occlusive |  |  |  |  |  | criteria met |
|  | disease |  |  |  |  |  |  |
| Pneumonia† | Aneurysmal | N/A | 6% (14/221) | NR | 30 days | NR | N/A |
|  | or occlusive |  |  |  |  |  |  |
|  | disease |  |  |  |  |  |  |
| Stroke† | Aneurysmal | N/A | 1% (3/221) | NR | 30 days | ≤ 2.9% | Acceptance |
|  | or occlusive |  |  |  |  |  | criteria met |
| Stroke‡ | disease | N/A | 0% (0/221) | NR | 30 days – 1 | ≤ 2.4% | Acceptance |
|  |  |  |  |  | year |  | criteria met |
| Anemia† | Aneurysmal | N/A | 0.5% (1/221) | NR | 30 days | NR | N/A |
|  | or occlusive |  |  |  |  |  |  |
|  | disease |  |  |  |  |  |  |
| Auricular | Aneurysmal | N/A | 1% (2/221) | NR | 30 days | ≤ 4.7% | Acceptance |
| fibrillation† | or occlusive |  |  |  |  |  | criteria met |
|  | disease |  |  |  |  |  |  |
| Abdominal | Aneurysmal | N/A | 0.5% (1/221) | NR | 30 days | NR | N/A |
| compartment | or occlusive |  |  |  |  |  |  |
| syndrome† | disease |  |  |  |  |  |  |
| Decubitus† | Aneurysmal | N/A | 0.5% (1/221) | NR | 30 days | NR | N/A |
|  | or occlusive |  |  |  |  |  |  |
|  | disease |  |  |  |  |  |  |
| Hemothorax† | Aneurysmal | N/A | 0.5% (1/221) | NR | 30 days | NR | N/A |
|  | or occlusive |  |  |  |  |  |  |
|  | disease |  |  |  |  |  |  |
| Small bowel | Aneurysmal | N/A | 0.5% (1/221) | NR | 30 days | NR | N/A |
| obstruction† | or occlusive |  |  |  |  |  |  |
|  | disease |  |  |  |  |  |  |
| Sepsis† | Aneurysmal | N/A | 0.5% (1/221) | NR | 30 days | NR | N/A |
|  | or occlusive |  |  |  |  |  |  |
|  | disease |  |  |  |  |  |  |
|  |  | N/A | 0% (0/221) | NR | 30 days | NR | N/A |

SOP 35-012 rev. B

| Summary of Safety and Clinical Performance | Page 30 of 42 |  |
| --- | --- | --- |
| AlboGraft™ Polyester Vascular Graft |  |  |

| Residual | Indication | Location | Device under | Device | Follow-up | Benchmark | Conclusions |
| --- | --- | --- | --- | --- | --- | --- | --- |
| risk in RMF |  |  | evaluation | under |  |  |  |
| and IFU |  |  | (literature, | evaluation |  |  |  |
|  |  |  | investigations, | (complaints/ |  |  |  |
|  |  |  | PMCF, | vigilance) |  |  |  |
|  |  |  | registries) |  |  |  |  |
| Aortoenteric | Aneurysmal | N/A | 0% (0/221) | NR | 30 days – 1 | NR | N/A |
| fistula‡ | or occlusive |  |  |  | year |  |  |
|  | disease |  |  |  |  |  |  |
| False | Aneurysmal | N/A | 0% (0/221) | NR | 30 days | NR | N/A |
| aneurysm‡ | or occlusive | N/A | 0% (0/221) | NR | 30 days – 1 | NR | N/A |
|  | disease |  |  |  | year |  |  |
| Lymphocele‡ | Aneurysmal | N/A | 4% (8/221) | NR | 30 days | NR | N/A |
|  | or occlusive |  |  |  |  |  |  |
|  | disease |  |  |  |  |  |  |
| Stenosis‡ | Aneurysmal | N/A | 0.5% (1/221) | NR | 30 day – 1 | NR | N/A |
|  | or occlusive |  |  |  | year |  |  |
|  | disease |  |  |  |  |  |  |
| Incisional | Aneurysmal | N/A | 1% (3/221) | NR | 30 days – 1 | NR | N/A |
| hernia† | or occlusive |  |  |  | year |  |  |
|  | disease |  |  |  |  |  |  |
| Other | Aneurysmal | N/A | 2% (4/221) | NR | 30 days – 1 | NR | N/A |
| (includes: | or occlusive |  |  |  | year |  |  |
| Incisional | disease |  |  |  |  |  |  |
| hernia, |  |  |  |  |  |  |  |
| hematoma, |  |  |  |  |  |  |  |
| lymphocele |  |  |  |  |  |  |  |
| and |  |  |  |  |  |  |  |
| neuropathic |  |  |  |  |  |  |  |
| pain at the |  |  |  |  |  |  |  |
| level of Scarpa |  |  |  |  |  |  |  |
| on the right |  |  |  |  |  |  |  |
| side) |  |  |  |  |  |  |  |
| NR= Not reported in the state of the art, no acceptance criterion was established |  |  |  |  |  |  |  |
| N/A= Not applicable, no comparison can be made due to no acceptance criterion |  |  |  |  |  |  |  |
| ** 99% of patients were treated for thoracic aortic aneurysms or acute or chronic dissections of the aorta. |  |  |  |  |  |  |  |
| *** 97% of patients were treated for occlusive disease. |  |  |  |  |  |  |  |
| †Postoperative procedure-related complication |  |  |  |  |  |  |  |
| ‡Postoperative device-related complication |  |  |  |  |  |  |  |

iv)	An overall summary of the clinical performance and safety

Analysis of Condition

The AlboGraft Polyester Vascular Graft is indicated for use in the replacement or repair of arteries affected with aneurysmal or occlusive disease, such as aortic aneurysms or peripheral vascular diseases. Aneurysms are bulges that develop as the vessel weakens, which can lead to ruptures and internal bleeding or the redirection of blood flow away from viral organs. Peripheral vascular occlusive disease

SOP 35-012 rev. B

| Summary of Safety and Clinical Performance | Page 31 of 42 |  |
| --- | --- | --- |
| AlboGraft™ Polyester Vascular Graft |  |  |

causes restricted blood flow, and most frequently affects the arteries of the lower limbs. If untreated, these conditions can lead to permanent adverse events (i.e., amputation) or death.

Current Treatment Options

Conservative and/or therapeutic treatments (i.e., exercise, medication, lifestyle modification) should be considered as initial treatment for peripheral vascular diseases. However, symptomatic cases and aneurysms of certain size or ruptured aneurysms should be repaired using endovascular (angioplasty and stenting) or open surgical methods (bypass with graft). Bypass surgery involves the surgical opening of a patient and replacement of diseased tissue with autologous vein grafts, other biological grafts (human umbilical veins, xenografts, or biosynthetic grafts), or synthetic grafts (PET or ePFTE). Synthetic grafts can be impregnated with compounds such as gelatin, albumin, and/or heparin to reduce permeability and clotting. While autologous veins are the recommended first option for bypass grafts, synthetic grafts provide an essential alternative when such veins are unavailable. Endovascular procedures are less invasive than open surgery and involve delivery of a stent to the disease site via percutaneous or surgical insertion of a catheter at a distal site. Data from the state-of-the-art evaluation suggests that endovascular repair results in better mortality rates than open surgery in the short-term, but long-term mortality reported >30 days was similar between the treatment groups. Long-term performance outcomes, such as patency, for endovascular repair were comparable to or worse than those observed for open surgical repair. Predominately, the rates of other safety outcomes were comparable between endovascular treatment and open surgical repair. Overall, the selection of revascularization technique is guided by the specific indication, the severity of the disease, and the availability of autologous conduits.

Benefit

The device under evaluation is intended for the replacement of diseased vessels, resulting from aneurysmal or occlusive disease. In 3 post-market studies with up to 2 years follow-up, 245 patients were treated with the device under evaluation for infrarenal abdominal aortic aneurysm, thoracic aortic aneurysm, acute or chronic Type A or B aortic dissection, atherosclerotic disease of the ascending aorta, peripheral arterial occlusive disease with aorto-iliac lesions, or iliac artery aneurysms. In one PMCF registry with 1 year follow-up, 221 patients were treated with the device under evaluation for aneurysmal or occlusive disease.

Primary patency (88.7-94.8%), secondary patency (91.9-95.1%), survival (89.6-99.0%), limb salvage (98.4%), and reintervention rates (2.9%) for the subject devices were comparable to similar treatments and within the acceptable limits reported in the state-of-the-art literature evaluation.

Risk

Risks inherent to device use are also associated with risk for the surgical procedure required to implant grafts. Risks identified in the IFU include bleeding, graft dilation, graft disintegration, graft Infection, graft stenosis, wound Infection, acute kidney failure, amputation, aneurysm, angina abdominalis, atrial fibrillation, bypass occlusions, chylothorax, compartment syndrome, death, deep vein thrombosis, descending aorta dissection, embolism, gluteal necrosis, groin seroma, hemiparesis, ileus, intraluminal graft thrombus, ischemia, myocardial infarction, paraparesis, paraplegia, paravalvular leak, perigraft air, pneumonia, pseudoaneurysm, renal dysfunction, respiratory failure, sternum instability, stroke, temporary

SOP 35-012 rev. B

| Summary of Safety and Clinical Performance | Page 32 of 42 |  |
| --- | --- | --- |
| AlboGraft™ Polyester Vascular Graft |  |  |

psychotic syndrome (including transient delirium), thrombosis, urinary tract infection, visceral ischemia, and wound hematoma.

None of the adverse events reported in the 3 post-market studies were directly attributed to the device under evaluation. However, the PMCF retrospective registry observed a total of 21 device-related complications in 18 patients throughout the 30-day follow-up period. The two most common device-related adverse events besides lymphocele (8/221, 4%) were graft infection (4/221, 2%) and graft thrombosis (4/221, 2%) requiring intervention. Graft infection (3%), graft thrombosis (3%) and superficial wound infections (3%) were the most frequent complication during follow-up between 30-days and 1 year. In this period, one death occurred due to graft infection. In addition, many of the adverse events were not reported in the clinical studies from the state of the art, and therefore there were no acceptance criteria established for those outcomes. Mortality, specific rates of myocardial infarction (e.g., cardiac complications) thrombosis, and amputation rates met the acceptance criteria established by the state of the art. No graft integrity outcomes (e.g., graft dilation, graft disintegration, etc.) were reported in the post-market studies or PMCF registry.

From 01 January 2019 to 31 May 2024, there were a total of 184 complaints associated with the device under evaluation and a total of 47,011 devices sold, resulting in an overall cumulative complaint rate of 0.391%. The complaint categories were red temperature indicator, defective temperature indicator, packaging issues, shipping damage, difficult to remove monofilament, hole in the graft, contamination by the user, barcode error, graft thrombosis/occlusion, leaking, service issue – incorrect product delivered, collagen fell off, difficult to remove Tyvek lid, and user error. There were no complaint trends based on device variant. However, there was an upward complaint trend over time for the Europe region. This was due to complaints of red temperature indicator and defective temperature indicator, for which CAPA 2024-010 has been opened and is currently in the investigation phase. During the same time frame, there were 9 MDRs with the top device problem codes of patient device interaction problem, device contaminated at the user facility, and material rupture. There were no MDRs involving death and 4 pertained to injury. The top patient problems was failure of implant (n=1, 11.1%).

Information about medicinal substances, animal or human tissues in the device, if any

The AlboGraft Polyester Vascular Grafts are impregnated with bovine collagen. The use of bovine collagen has been assessed by the following:

Nonclinical testing evaluating the tissue reaction and resorption profile of the Albograft device in comparison to the Hemashield (predicate device manufactured by Boston Scientific Inc.) in an in-vivo sheep model over 0–6-week time period and to assess systemic toxicity at 6 weeks (Section Error! R eference source not found.).

A separate state of the art report evaluating the use of bovine Collagen (Error! Reference source not f ound.).

Updates to the current systematic literature reviews for both the state of the art and the device under evaluation demonstrated comparable or better performance of grafts impregnated with bovine collagen to treatment alternatives, without any differences in safety outcomes.

Risk mitigations

SOP 35-012 rev. B

| Summary of Safety and Clinical Performance | Page 33 of 42 |  |
| --- | --- | --- |
| AlboGraft™ Polyester Vascular Graft |  |  |

Risk mitigation has been performed as described in the risk management documentation using the following risk control options:

- inherent safety by design;

- protective measures in the medical device itself or in the manufacturing process;

- information for safety.

It should be noted that the IFU specifies that the AlboGraft Polyester Vascular Graft should only be used by vascular surgeons. Considering how complex a surgical implantation of vascular grafts is and how many factors can influence the patient’s condition, it is left to the surgeon to define suture technique and graft type as well as the therapy to adopt before, during and after the operation.

PMS, including PMCF studies, is conducted to monitor the safety and performance of the device throughout its lifetime.

These mitigations are considered sufficient to mitigate risks associated with use of the AlboGraft Polyester Vascular Graft.

Conclusions

Based on this clinical evaluation, which includes non-clinical and clinical data, there is sufficient data to demonstrate conformity to the applicable requirements and confirm the that the AlboGraft Polyester Vascular Graft is safe and performs as intended and claimed by LeMaitre Vascular, Inc. The AlboGraft Polyester Vascular Graft is a state-of-the-art device for use in the replacement or repair of arteries affected with aneurysmal or occlusive disease, such as such as infrarenal abdominal aortic aneurysm or aorto-iliac occlusive lesions. Review of the post-market data, information materials, and the risk management documentation provided by LeMaitre Vascular, Inc confirms that the risks are appropriately identified and consistent with the state of the art, and that the risks associated with the use of the device are acceptable when weighed against the benefits.

v)	Ongoing or planned post-market clinical follow-up

The manufacturer conducts ongoing post-market surveillance (PMS) of the subject device according to the following procedure, SOP28-001. Post-Market Clinical Follow-up (PMCF) activities are planned for the subject device. A multi-stepped approach will be used to substantiate the performance claims of the device and ensure that the risk/benefit remains positive. First, a thorough literature review was conducted to capture all relevant and up to date published information regarding the Albograft device. The second step will involve completion of multi-center study in Europe. Contract negotiations were completed in Q2 of 2022, with the study start ongoing.

The purpose of the study is to conduct a retrospective analysis on the performance and safety of the AlboGraft Vascular Graft on patients undergoing surgical treatment for aneurysmal or occlusive disease with a maximal follow up of one year.

SOP 35-012 rev. B

| Summary of Safety and Clinical Performance | Page 34 of 42 |  |
| --- | --- | --- |
| AlboGraft™ Polyester Vascular Graft |  |  |

It is anticipated this study will be extended into an on-going registry to confirm the safety and effectiveness throughout the expected lifetime of the device through the proactive and continuous collection of data.

- Possible diagnostic or therapeutic alternatives:

  - Peripheral Vascular Repair: Invasive treatments are not recommended for asymptomatic peripheral arterial disease. In many cases, intermittent claudication caused by peripheral arterial disease can be managed with medical treatment (e.g., smoking cessation interventions, statin therapy, antiplatelet therapy) or exercise therapy. However, SVS recommends invasive (endovascular or surgical) treatment for patients with “significant functional or lifestyle-limiting disability when there is a reasonable likelihood of symptomatic improvement with treatment, when

pharmacologic or exercise therapy, or both, have failed, and when the benefits of treatment outweigh the potential risks.”6 Invasive treatment should be individualized to the patient. For instance, endovascular procedures are recommended over open surgery for focal occlusive disease of the superficial femoral artery, whereas surgical bypass is recommended as an initial revascularization strategy for patients with

diffuse femoro-popliteal disease or extensive calcification of the superficial femoral artery (depending on patient anatomy).7 ESC/ESVS suggest endovascular therapy as the first choice of treatment for femoro-popliteal lesions <25 cm and surgical bypass (especially when using the great saphenous vein) for occlusion/stenosis >25 cm in length.

Bypass may be achieved using autologous vein, biological graft such as human umbilical vein, synthetic grafts (typically ePTFE [also referred to as PTFE] or Dacron), or biosynthetic grafts (e.g., LeMaitre Omniflow II, which is constructed of polyester mesh and ovine collagen). Heparin-bonded synthetic grafts, designed to reduce risk of thrombosis, have also been introduced to the market. The consensus by professional societies, including the European Society of Cardiology and European Society for Vascular Surgery, is that autologous vein should be used for bypass whenever possible, but the use of a prosthetic graft should be considered in the absence of suitable vein.6,7 The clinical practice guidelines do not contraindicate the use of synthetic grafts in the coronary arteries, and prosthetic grafts are required for coronary artery bypass grafting when the availability of suitable autologous conduits is limited8. However, the nonsystematic review by Desai et al. (2011) concluded that existing synthetic grafts do not meet the equivalent function and durability of the internal mammary artery or long saphenous vein in coronary artery bypass grafting8. Therefore, a contraindication of the use of grafts like AlboGraft in the coronary arteries is appropriate.

-- Abdominal aortic aneurysm repair: Endovascular repair for AAAs became available in 1991. While endovascular interventions are increasing in use, open

SOP 35-012 rev. B

| Summary of Safety and Clinical Performance | Page 35 of 42 |  |
| --- | --- | --- |
| AlboGraft™ Polyester Vascular Graft |  |  |

repair remains the standard procedure for AAA repair. Given that there are no proven medical therapies available to slow the expansion of AAAs, surgical interventions are typically required when the growth exceeds a certain threshold (≥5.5 cm for men and ≥5.0cm for women) or there is a rupture. When rapid AAA growth is observed (≥1 cm/year) or there is an increase in symptoms, more urgent referral to a vascular surgeon is recommended.6 Open surgical repair involves a large incision, along the abdomen in the case of AAA, removal of the damaged vessel at the site of the aneurysm and implantation of a graft to replace that segment. Polyethylene terephthalate, also known by its brand name Dacron, is the most frequently used material in open surgical repair of AAA for the last 60 years. Dacron grafts are available with different kinds of impregnation (i.e. gelatin, albumin, etc) to decrease the porosity of the graft. Expanded polytetrafluoroethylene (PTFE) is an alternative synthetic graft material. Endovascular aneurysm repair is a minimally invasive option involving a smaller incision in the groin and the insertion of stent grafts via catheter, via either percutaneous or surgical access, in the artery that is then threaded up to the location of the aneurysm. The placement of the stent graft then acts to support the aneurysm. Unlike grafts used in open repair, a stent graft is meant to seal the sac from the inside of the aneurysm, while the aneurysm wall is left untouched. Most of the stent graft devices require a degree of oversizing of the graft relative to the vessel (≈10-25%) to ensure adequate sealing and fixation. Percutaneous endovascular stent placement is associated with fewer access-related complications, such as groin infection and lymphocele.

- Suggested profile and training for users

The AlboGraft Vascular Graft is an implant intended for use by experienced vascular surgeons trained in the procedures for which they are intended.

| 8.0 | Reference to any harmonized standards and CS applied |  |
| --- | --- | --- |
|  | Standard Title | Standard Reference: |
|  |  | Revision Year |
|  | Sterilization of medical devices. Requirements for medical devices to be | EN 556-2:2015 |
|  | designated “STERILE”. Part 2: Requirements for aseptically processed |  |
|  | medical devices |  |
|  | Information supplied by the manufacturer of medical devices | EN 1041:2008 |
|  | Cardiovascular implants and extracorporeal systems – Vascular prostheses -- | ISO 7198:2016 |
|  | Tubular vascular grafts and vascular patches |  |
|  | Biological evaluation of medical devices – Part 1: Evaluation and testing | ISO 10993-1:2009 |
|  | Biological evaluation of medical devices – Part 3: Tests for genotoxicity, | ISO 10993-3:2009 |
|  | carcinogenicity and reproductive toxicity |  |
|  | Biological evaluation of medical devices – Part 4: Selection of tests for | EN ISO 10993-4:2006 |
|  | interactions with blood |  |
|  | Biological evaluation of medical devices – Part 5: Tests for in vitro | ISO 10993-5:2009 |
|  | cytotoxicity |  |

SOP 35-012 rev. B

| Summary of Safety and Clinical Performance | Page 36 of 42 |  |
| --- | --- | --- |
| AlboGraft™ Polyester Vascular Graft |  |  |

| Biological evaluation of medical devices – Part 6: Tests for local effects after | EN ISO 10993-6:2007 |
| --- | --- |
| implantation |  |
| Biological evaluation of medical devices – Part 10: Tests for irritation and | ISO 10993-10:2010 |
| delayed-type hypersensitivity |  |
| Biological evaluation of medical devices – Part 11: Tests for systemic | ISO 10993-11:2018 |
| toxicity |  |
| Biological evaluation of medical devices Part 17: Establishment of allowable | EN ISO 10993-17:2008 |
| limits for leachable substances |  |
| Packaging for terminally sterilized medical devices – Part 1: Requirements | ISO 11607-1:2006 |
| for materials, sterile barrier systems and packaging systems |  |
| Packaging for terminally sterilized medical devices – Part 2: Validation | ISO 11607-2:2006 |
| requirements for forming, sealing and assembly processes |  |
| Sterilization of medical devices – Microbiological methods – Part 1: | ISO 11737-1:2006 |
| Determination of a population of microorganisms on products |  |
| Tests of sterility performed in the definition, validation and maintenance of a | ISO 11737-2:2009 |
| sterilization process |  |
| Aseptic processing of health care products – Part 1: General requirements | ISO 13408-1:2008 |
| Medical devices – Quality management systems – Requirements for | EN ISO 13485:2016 |
| regulatory purposes |  |
| Sterilization of health care products – Liquid chemical sterilizing agents for | ISO 14160:2011 |
| single-use medical devices utilizing animal tissues and their derivatives – |  |
| Requirements for characterization, development, validation and routine |  |
| control of a sterilization process for medical devices |  |
| Cleanrooms and associated controlled environments – Part 1: Classification | ISO 14644-1:2015 |
| of air cleanliness |  |
| Medical devices – Application of risk management to medical devices | EN ISO 14971:2012 |
| Medical devices — Symbols to be used with medical device labels, labelling | EN ISO 15223-1:2016 |
| and information to be supplied —Part 1: General requirements |  |
| Medical devices utilizing animal tissues and their derivatives – Part 1: | ISO 22442-1:2015 |
| Application of risk management |  |
| Medical devices utilizing animal tissues and their derivatives – Part 2: | ISO 22442-2:2015 |
| Controls on sourcing, collection and handling |  |
| Medical devices utilizing animal tissues and their derivatives – Part 3: | ISO 22442-3:2007 |
| Validation of the elimination and/or inactivation of viruses and TSE agents |  |

Bibliography:

- van Det RJ, Vriens BH, van der Palen J, Geelkerken RH. Dacron or ePTFE for femoro-popliteal above-knee bypass grafting: short- and long-term results of a multicentre randomised trial.

European journal of vascular and endovascular surgery : the official journal of the European Society for Vascular Surgery. 2009;37(4):457-463.

- Hsu RB, Chen JW. Low incidence of late pseudoaneurysm and reoperation after conventional repair of acute type a

aortic dissection. Journal of cardiac surgery. 2014;29(5):641-646.

- Rajbanshi BG, Joshi D, Pradhan S, et al. Primary surgical repair of coarctation of the aorta in adolescents and adults: Intermediate results and consequences of hypertension. European Journal of Cardio-thoracic Surgery. 2019;55(2):323-330.

- Tamura A, Yamaguchi A, Yuri K, et al. Clinical experience with a new vascular graft free from biodegradable material.

Interactive cardiovascular and thoracic surgery. 2011;12(5):758-761.

- Lamelas J, Chen PC, Loor G, LaPietra A. Successful Use of Sternal-Sparing Minimally Invasive Surgery for Proximal Ascending Aortic Pathology. Annals of Thoracic Surgery 2018;106(3):742-748.

SOP 35-012 rev. B

| Summary of Safety and Clinical Performance | Page 37 of 42 |  |
| --- | --- | --- |
| AlboGraft™ Polyester Vascular Graft |  |  |

  - Society for Vascular Surgery Lower Extremity Guidelines Writing G, Conte MS, Pomposelli FB, et al. Society for Vascular Surgery practice guidelines for atherosclerotic occlusive disease of the lower extremities: management of asymptomatic disease and claudication. J Vasc Surg. 2015;61 (3 Suppl):2S-41S.

  - Aboyans V, Ricco JB, Bartelink MEL, et al. Editor's Choice - 2017 ESC Guidelines on the Diagnosis and Treatment of Peripheral Arterial Diseases, in collaboration with the European Society for Vascular Surgery (ESVS). European journal of vascular and endovascular surgery: the official journal of the European Society for Vascular Surgery.

2018;55(3):305-368.

  - Desai M, Seifalian AM, Hamilton G. Role of prosthetic conduits in coronary artery bypass grafting. Eur J Cardiothorac Surg. 2011;40(2):394-398.

- Revision History

| SSCP | Date issued | Change description | Revision validated by the NotifiedBody |  |
| --- | --- | --- | --- | --- |
| revision |  |  |  |  |
| number |  |  |  |  |
| A | 24/03/2022 | Initial release | ☐ | Yes; Validation language: English |
|  |  |  | (only applicable for class IIa or some IIb implantable devices |  |
|  |  |  | (MDR, Article 52 (4) 2nd paragraph) for which the SSCP is not yet |  |
|  |  |  | validated by the NB) |  |
|  |  |  | ☒ | No, pending initial review |
| B | 09/12/2022 | Added section 10 | ☐ | Yes; Validation language: English |
|  |  | patient info, and made | (only applicable for class IIa or some IIb implantable devices |  |
|  |  | updates per BSI | (MDR, Article 52 (4) 2nd paragraph) for which the SSCP is not yet |  |
|  |  | feedback | validated by the NB) |  |
|  |  |  | ☒ | No, pending initial review |
| C | 22/05/2024 | Updated device | ☒ | Yes; Validation language: English |
|  |  | description and model | (only applicable for class IIa or some IIb implantable devices |  |
|  |  | numbers | (MDR, Article 52 (4) 2nd paragraph) for which the SSCP is not yet |  |
|  |  |  | validated by the NB) |  |
|  |  |  | ☐ | No |
| D | 11/11/2024 | Periodic update | ☐ | Yes; Validation language: English |
|  |  |  | (only applicable for class IIa or some IIb implantable devices |  |
|  |  |  | (MDR, Article 52 (4) 2nd paragraph) for which the SSCP is not yet |  |
|  |  |  | validated by the NB) |  |
|  |  |  | ☒ | No, NB approval not needed. The safety and |
|  |  |  | performance of the device has not changed since the last NB |  |
|  |  |  | approval. |  |

SOP 35-012 rev. B

| Summary of Safety and Clinical Performance | Page 38 of 42 |  |
| --- | --- | --- |
| AlboGraft™ Polyester Vascular Graft |  |  |

10. Patient information

A summary of the safety and clinical performance of the device, intended for patients, is given below.

This Summary of Safety and Clinical Performance (SSCP) is intended to provide public access to an updated summary of the main aspects of the safety and clinical performance of the device. The information presented below is intended for patients or lay persons. Your healthcare provider has a more extensive summary of safety and clinical performance.

The SSCP is not intended to give general advice on the treatment of a medical condition. Please contact your healthcare professional in case you have questions about your medical condition or about the use of the device in your situation. This SSCP is not intended to replace an implant card or the instructions for use to provide information on the safe use of the device.

- Device general information

  - Device trade name

    - AlboGraft Polyester Vascular Graft (Graft)

  - Producer; name and address

    - LeMaitre Vascular, Inc. 32 Third Avenue, Burlington, MA 01803

  - Basic UDI-DI

    - 08406631AlboGraftR6

  - Year when the device was first CE-marked

    - 2011

- Intended use of the device

  - Intended purpose

    - The Grafts are intended to be used as a replacement for diseased blood vessels.

  - Indications and intended patient groups

    - The Graft is used for repair or replacement of diseased blood vessels or during surgery to correct blood flow.

    - The product is designed for adults of any gender or ethnicity in need of repair or replacement of diseased blood vessels or during surgery to correct blood flow.

  - Do not use for:

    - The Grafts should not be used in coronary arteries and in patients with known or suspected negative reactions to any form of protein derived from cows.

- Device description

  - Device description and material/substances in contact with patient tissues

    - The Grafts are made of synthetic material designed to replace sections of damaged or diseased arteries. They are made of polyester PET thread woven into a continuous tube. The following materials are what make up the graft, polyester cow collagen and a preserving agent. All materials have passed testing to ensure they are safe to use

  - Information about medicinal substances in the device, if any

    - n/a

  - Description of how the device is achieving its intended mode of action

    - Per regulations, the Graft achieves its effect through non-medicinal means. It achieves this goal as a physical barrier device as its mode of action.

SOP 35-012 rev. B

| Summary of Safety and Clinical Performance | Page 39 of 42 |  |
| --- | --- | --- |
| AlboGraft™ Polyester Vascular Graft |  |  |

  - Description of accessories, if any i. n/a

- Risks and warnings

Contact your healthcare professional if you believe that you are experiencing side effects related to the device or its use or if you are concerned about risks. This document is not intended to replace a consultation with your healthcare professional if needed.

| Patient-Related Adverse Event | Severity | Occurrence | RPN |
| --- | --- | --- | --- |
| Stroke | 8 | 2 | 16 |
| Complete/partial paralysis in legs or | 8 | 2 | 16 |
| lower abdomen |  |  |  |
| Partial paralysis of both legs | 8 | 2 | 16 |
| Heart attack | 8 | 2 | 16 |
| Kidney stops working | 8 | 2 | 16 |
| Blood flow limited | 8 | 2 | 16 |
| Blood clot that blocks blood flow | 8 | 2 | 16 |
| Blood clotting in vein | 8 | 1 | 8 |
| The process of losing blood from the | 8 | 2 | 16 |
| body |  |  |  |
| Growth of germs in or around the wound | 8 | 2 | 16 |
| Bulge in the wall of an artery | 8 | 2 | 16 |
| Lung inflammation with tiny fluid filled | 8 | 1 | 8 |
| air sacs |  |  |  |
| Loss or removal of a body part | 8 | 1 | 8 |
| Death | 10 | 1 | 10 |
| Injured blood vessel wall that leads to | 8 | 2 | 16 |
| leaking |  |  |  |
| The formation of a blood clot inside the | 8 | 1 | 8 |
| artery and vein |  |  |  |
| A serious condition that makes it difficult | 8 | 1 | 8 |
| to breathe on your own |  |  |  |
| An irregular and often very rapid heart | 8 | 1 | 8 |
| rhythm that can lead to blood clots in the |  |  |  |
| heart |  |  |  |
| Fluid leaking into space between lung | 8 | 2 | 16 |
| and chest wall |  |  |  |
| Temporarily not in ones right mind | 8 | 1 | 8 |
| Total or partial paralysis of one side of | 8 | 1 | 8 |
| the body |  |  |  |
| Poor blood supply to nervous system | 8 | 1 | 8 |
| Numbness running down buttocks to | 8 | 1 | 8 |
| back of legs |  |  |  |
| Intestines not able to move food through | 8 | 1 | 8 |
| digestive system |  |  |  |
| Severe and sudden kidney failure | 8 | 1 | 8 |
| Swelling or bleeding occurs within a | 8 | 1 | 8 |
| compartment usually in legs, feet, arms |  |  |  |
| or hands |  |  |  |
| Leaking caused by space left between | 8 | 1 | 8 |
| heart and valves |  |  |  |
| Tear in inner layer of your aorta, the | 8 | 1 | 8 |
| main artery that delivers blood you’re |  |  |  |
| your heart to your body |  |  |  |
| Blood clot develops in veins deep in your | 8 | 1 | 8 |
| body |  |  |  |

SOP 35-012 rev. B

| Summary of Safety and Clinical Performance | Page 40 of 42 |  |
| --- | --- | --- |
| AlboGraft™ Polyester Vascular Graft |  |  |

| Excessive movement due to disruption of | 6 | 1 | 6 |
| --- | --- | --- | --- |
| the wires connecting the surgically |  |  |  |
| divided sternum |  |  |  |
| Patient will need a two-part surgery | 8 | 2 | 16 |
| combining open microsurgery and |  |  |  |
| endovascular coiling. The reason for this |  |  |  |
| surgery is to coil the entire diseased part |  |  |  |
| of the blood vessel and then bypass the |  |  |  |
| blood flow to the specific location in the |  |  |  |
| brain |  |  |  |
| A bad bruise | 6 | 1 | 6 |
| A collection of fluid that builds up under | 6 | 1 | 6 |
| the surface of your skin |  |  |  |
| An infection in any part of your urinary | 8 | 1 | 8 |
| system |  |  |  |
| Local pain in the stomach area | 8 | 1 | 8 |
| Device-Related Adverse Event | Severity | Occurrence | RPN |
| The graft is stretched or enlarged beyond | 8 | 2 | 16 |
| normal |  |  |  |
| Graft losing cohesion or strength | 8 | 2 | 16 |
| Graft is narrowing | 8 | 1 | 8 |
| Graft is infected | 8 | 1 | 8 |
| Air in or around the graft | 8 | 1 | 8 |

- How potential risks have been controlled or managed

    - Analysis have concluded that the benefits outweigh any residual risks and that the risk has been reduced as far as possible

- Remaining risks and undesirable effects

    - Please refer to the device IFU or your healthcare provider.

- Warnings and precautions

  - Your new device is a foreign body and therefore needs close monitoring and careful observation. It may take 6-8 weeks for full recovery.

  - After placement, the implant area maybe swollen and tender for up to a week.

  - Observe for any new redness or tenderness

  - Observe for any opening in the incisions.

  - Observe for numbness, tingling or pain in the leg.

NOTE if you experience any of the above (2-5) please contact your provider.

- Do not puncture or manipulate the graft.

- You may shower according to your provider’s instructions.

- Swelling in the extremity is expected because of increased blood flow. Move according to your provider’s instructions, if the graft was implanted in your leg. Keep your leg elevated above your heart.

- It is preferable to have the graft covered for the first week to protect skin and incisions. (Follow your provider’s instructions).

- Keep bandages or compression bandages on as per your provider’s instructions.

- If your staples have been removed, you will probably have Steri-Strips (small pieces of tape) across your incision. Wear loose clothing that does not rub against your incision.

- You may shower or get the incision wet, once your provider says you can. DO NOT soak, scrub, or have the shower beat directly on them. If you have Steri-Strips, they will curl up and fall off on their own after a week.

- DO NOT soak in the bathtub, a hot tub, or swimming pool. Ask your provider when you can start doing these activities again.

SOP 35-012 rev. B

| Summary of Safety and Clinical Performance | Page 41 of 42 |  |
| --- | --- | --- |
| AlboGraft™ Polyester Vascular Graft |  |  |

  - Your provider will tell you how often to change your dressing (bandage) and when you may stop using one. Keep your wound dry. If your incision goes to your groin, keep a dry gauze pad over it to keep it dry.

  - Clean your incision with soap and water every day once your provider says you can. Look carefully for any changes. Gently pat it dry.

  - DO NOT put any lotion, cream, or herbal remedy on your wound without asking your provider first if that is ok.

  - Bypass surgery does not cure the cause of the blockage in your arteries. Your arteries may become narrow again.

  - Eat a heart-healthy diet, exercise, stop smoking (if you smoke), and reduce stress. Doing these things will help lower your chances of developing a blocked artery again.

  - Your provider may give you medicine to help lower your cholesterol.

  - If you are taking prescriptions for high blood pressure or diabetes, take them as prescribed.

  - Your provider may ask you to take aspirin or a medicine called clopidogrel (Plavix) when you go home. These medicines keep your blood from forming clots in your arteries. DO NOT stop taking them without talking to your provider first.

- Summary of clinical evaluation and post-market clinical follow-up

    - Clinical background of the device

The Graft is categorized as Class II device in US and Class III device in EU. The Grafts are made of synthetic material and designed to replace sections of damaged or malfunctioning arteries. They are made of polyester thread woven into a seamless tube. In response to a range of surgical indications, the Grafts are offered in two designs: double velour knitted fabric and double velour woven fabric. The knitted grafts are designed with a run-proof structure to reduce the risk of fraying or wearing down at their ends. The velour grafts have low profile loops on their endoluminal surface to avoid any lumen reduction, and high profile loops on their outer surface to promote graft anchoring into the surrounding tissues. The AlboGraft will also be available with or without collagen coating. All of the grafts are crimped in parallel rings so that their tubular shape is maintained without kinking.

The Grafts are available with removable external spiral reinforcement made of a thread, allowing for easy identification with x-ray. The external spiral reinforcement is removable, helping the joingin of the vessel to the graft.

The Grafts are made with collagen to reduce leakage so that no pre-clotting is necessary. The process of using bovine collagen maintains both the original structure of the material, and the structural characteristics of the graft, i.e. flexibility and softness.

    - The clinical evidence for the CE-marking

The device was first approved for CE mark under LeMaitre Vascular Inc. in 2011. Studies were conducted to ensure the grafts were safe and effective. See the IFU for further details.

    - Safety

There are ongoing clinical trials on this graft that will be used to confirm the safety and performance throughout the expected lifetime of the device through the proactive and continuous collection of data.

SOP 35-012 rev. B

| Summary of Safety and Clinical Performance | Page 42 of 42 |  |
| --- | --- | --- |
| AlboGraft™ Polyester Vascular Graft |  |  |

- Possible alternatives

When considering alternative treatments, it is recommended to contact your healthcare professional who can take into account your individual situation.

- Suggested training for users

  - This device is intended to be used by surgeons. Considering how complex this surgery is, it is left to the surgeon to proper surgery and graft type as well as the therapy to adopt before, during and after the operation.

SOP 35-012 rev. B